## Role 2 Technical Requirements
Lisa 1
|  | The tenderer must indicate the maximum delivery time of the system and the number of systems to be produced in parallel. | The delivery time of the first system is 10 months from the order; - The delivery time of the second system is 15 months from the order; - The delivery time of each next system is + 5 months to the previous one. If the order is submitted at the same time. 2 systems are produced in parallel in different phases.
|  | The tenderer should provide country of manufacturing*. | ESTONIA | * informative and not binding on the contract
Requirement number | Requirement Value MUST or SHALL | Requirement description | Development Status: The tenderer must insert one of the following terms: * EX (Existing Feature) – the system supplier's proposed solution exists already; * MD (Modification Required) – the system supplier's existing solution requires modification in order to meet the specified requirement; * DV (Development Required) – the system supplier's proposed solution does not yet exist. | Remarks: The tenderer's own descriptions and solutions for the contracting authority's requirements (mandatory on requirements which are negotiable (requirements with the value 'SHALL'))
ABOUT
Mobile Modular Hospital (MMH) solution (Role 2) is a container solution for a field hospital of military medical function which is planned to provide medical support during training excercises as well as in emergency and crisis situations. Role 2 solution consists of containers and tents with different functions – such as laboratory containers, triage tents, surgery containers, emergency medicine containers, intensive care containers, medical treatment containers, "ward" tents, sterilizer container, pharmacy container, logistics container, sanitary container, resource (electricity – water, electricity – oxygen) containers, that can be interconnected by connecting tents and walkways, if necessary, it has to be possible to add additional containers to the solutuon, according to the contracting authority's needs. Role 2 solution must be usable in locations without the possibility of connecting to utility networks. Role 2 solution is notionally allocated into two parts (legs) which are meant to function both together and separately. If the field hospital is in motion, the left side of the field hospital shall be the first to move. The deployment and repacking of the solution has to be possible within 1 hour at most, with a fully assembled team. The fully assembled team consists of: 3 persons in command, 5 persons per unloaded container (surgery, emergency medicine, intensive care, laboratory, connection container, ward container, pharmacy, sterilizer) and 2 persons for other containers (resources, medical logistics) – altogether 74 persons.
DETAILED DESCRIPTION OF THE FUNCTION
The field hospital offers medical care for the wounded and ill personnel of the Infantry Brigade at a NATO ROLE 2 level. In the triage tent, patients will be sorted in accordance with their medical care needs and their urgency; life-saving first aid shall be given if necessary. With the use of a connection container, other containers are connected into a single entity, first and foremost the triage, emergency medicine, surgery and intensive care container. The interconnection container in the laboratory container configuration will be supporting other containers' work, by offering diagnostics support, storage space, as well as hand- and instrument-washing capabilities. The emergency medicine container is for resuscitating patients and preparing them for surgery or intensive care or medical care facility or back to their unit if they have less serious injuries. The container must enable handling two patients simultaneously. The surgery facility is for operating on patients at a primary surgery level, by using the principles of controlling surgery (preservation of life, limb or organ functionality is primary). Up to two patients will be operated on simultaneously. In the intensive care container, up to four patients are cared for who require high-level medical care, by stabilizing them for transport or relocation to the medical treatment container. In the medical treatment container, patients who are ill, wounded or suffer from post-traumatic stress disorder will be treated, whose prognosis enables them to return to their unit within a few days. In the medical treatment container, NATO ROLE 1 level medical care will be possible to administer. Also, patients who require less medical care and who can be transported with a regular evacuation vehicle will be waiting for evacuation in the medical treatment container. The medical treatment container requires 20 bed spaces. The sterilization container is used for sterilizing surgical instruments for the field hospital and the nearby aid stations. The medical logistics container and storage container are used to provide pharmacological services for the field hospital and its nearby aid stations; they are equipped with medicine and other medical consumables and blood derivatives. The resource container produces electricity for the field hospital if necessary, but also water for consumption that has drinking water quality and medical oxygen. Oxygen is also stored in oxygen tanks, in order to supply aid stations and medical evacuation utilities which operate around the field hospital. For production of electricity, the emergency medicine, surgery and medical logistics containers are autonomous for the event that electricity cannot be obtained from the resource container or any other source. The emergency medicine and surgery container are autonomous for production of oxygen in the event that medical oxygen cannot be obtained from the resource container. The tenderer must manufacture the entire Role 2 solution in accordance with STANAG 2228, 2060, 2136, 2558, 2535, 2231 and 2481. The tenderer is allowed to submit a technical description which proves compliance with the aforementioned STANAG's. The tenderer must ensure that the Role 2 system and its different elements function in a way that a reasonable person would understand their functionality.
1. GENERAL
1.1. Capacity Indicators of MMH Role 2 Solution
1.1.1. | MUST | Role 2 two leg capacity: 17 15 containers and 6 high-pressure tents, 19 high-pressure corridors | EX
1.1.2. | MUST | Containers and tents are interconnected by tents and corridors. | EX
1.1.3. | MUST | It must be possible to use and connect the containers and tents in different combinations. | EX
1.1.4. | MUST | Container must have an expandable two-in-one solution. A two-in-one container means that one side of the container can be extended and a working space with the size of two containers can be created. This must be performed easily and without additional accessories. | EX | Two-in-one container is designed so that one side of the container can be expanded easily and without additional accessories. Container is foldable type, opening is combined with integrated hydraulic system and gas-springs.The opening of the expandable floor of the 2:1 containers is performed electro-hydraulically and the roof manually by gas springs. The necessary power supply is drawn from a battery bank integrated into the container.
1.1.5. | MUST | Container must have an expandable three-in-one solution. A three-in-one container means that both sides of the container can be extended and a working space with the size of three containers can be created. This must be performed easily and without additional accessories. | EX | Three-in-one container is designed so that both sides of the container can be expanded easily and without additional accessories. Container is foldable type, opening is combined with integrated hydraulic system and gas-springs.The opening of the expandable floor of the 3:1 containers is performed electro-hydraulically and the roof manually by gas springs. The necessary power supply is drawn from a battery bank integrated into the container.
1.1.6. |  | List and numbers of containers, tents and corridors of Role 2 two-part (2 legs) solution | EX
The following is a description of the set by units; below are the units' detailed descriptions.
1.1.6.1 | MUST | Laboratory (regular container) – 2 Connection (regular container) – 1 Tent “triage” – 2 High-pressure corridors – 11 | EX
1.1.6.2 | MUST | Surgery (three-in-one container) – 2 | EX
1.1.6.3 | MUST | Emergency medicine (three-in-one container) – 2 | EX
1.1.6.4 | MUST | Intensive care (three-in-one container) – 2 High-pressure corridors – 2 | EX
1.1.6.5 | MUST | Medical treatment (two-in-one container) – 2 Tent “ward” – 4 High-pressure corridors – 6 | EX
1.1.6.6 | MUST | Sterilizer (three-in-one container) – 1 | EX
1.1.6.7 | MUST | Pharmacy (two-in-one container) – 1 | EX
1.1.6.8 | MUST | Logistics container (regular container) – 1 | EX
1.1.6.9 | MUST | Resource 1 (electricity, oxygen) (regular container) – 1 | EX
1.1.6.10 | MUST | Resource 2 (electricity, water) (regular container) – 1 | EX
1.1.6.11 | MUST | Sanitary container (regular container) – 1 | EX
1.1.7. |  | Measurements of containers and tents | EX
1.1.7.1 | MUST | Container measurements during transport and regular container measurements (ISO 1CC 20’): | EX
|  | length – 6058 mm; height – 2591 mm; width – 2438 mm; max gross weight – 10,000 kg. | EX
1.1.7.2. | MUST | Regular container measurements in the working position: | EX
|  | Floor area (working area) – in the laboratory and connection container at least 10 m2 and 9,5 m2 in the logistics container. Interior height – at least 2200 mm | EX
1.1.7.3 | MUST | Two-in-one container measurements in the working position: | EX
|  | Floor area (working area) – at least 17 m2, if the container has a technical room (pharmacy) or at least 19 m2, if the container has no technical room (ward). Interior height – at least 2200 mm in the middle section and at least 2000 mm in the extended section. | EX
1.1.7.4 | MUST | Three-in-one container measurements in the working position: | EX
|  | Floor area (working area) – at least 16 m2, if the container has a technical room (surgery, emergency medicine) or at least 25 m2, if the container has no technical room (intensive care, sterilization). Interior height – at least 2200 mm in the middle section and at least 2000 mm in the extended section. | EX
1.1.7.5 | MUST | High-pressure tent measurements in the working position: | EX
|  | length – 10,000 mm (+/- 10%); height – 3100 mm (+10%); width – 6000 mm (+10%). | EX
1.2. General Requirements of MMH Role 2 Solution
1.2.1. | MUST | The deployment and repacking of the solution has to be possible within 1 hour at most with a fully-assembled team. | EX
1.2.3. | MUST | The constructional and functional solutions of the container must allow for the simple, safe and fast operation of the container. | EX
1.2.4. | MUST | The locations, placement and fastenings of equipment must facilitate its preservation and maintain its working order. The container has to include integrated equipment as much as possible. Equipment integration means that the equipment is fully attached to the structural parts of the container and is optimized for its intended use. The integrated equipment must be ready for immediate use after the container is set to the working position. It must be possible to remove the integrated equipment from the structural elements if necessary. | EX
1.2.5. | MUST | No container can have protruding parts (chimneys, antennas, lights etc.), which extend outwards from the container's general gauges. The containers must be possible to pack next to and on top of one another without any additional work. | EX
1.2.6. | MUST | The container must comply with the requirements in this document throughout its operational life. The operational life of the container must be at least 15 years and the operational life of the technical devices must be at least 5 years, in accordance with the conditions set in clause 1.3. The tenderer shall give their own estimation of the operational life of the system. | EX | Planned operational lifetime of the containers is 20 years, for equipment and devices, planned operational life time is 10 years
1.2.7. | MUST | Container walls must have manuals in English for the use of the system and its devices, exact locations and level of detail for its information shall be agreed upon during the design stage. | EX
1.2.8. | MUST | The tenderer is obligated to submit the information required for the identification/codification of the containers to a competent codification authority. The NSN code must be issued by the tenderer by the time of delivery of the R2 at the latest. | MD
1.2.9. | MUST | The NSN code must be permanently and clearly marked on the outside of containers. | MD
1.2.10. | MUST | The delivered solution must include two paper copies and a printable PDF of the user instructions and user maintenance and repair instructions in Estonian or Lithuanian (as per the buyer) and English. The main design is delivered digitally (as .dwg and .pdf files). All medical equipment must include maintenance guides, as well as the necessary user names and passwords for entry into the technical menus. | EX
1.2.11. | MUST | The tenderer must provide free user training and user maintenance training courses after the delivery of the solution. Those who passed the maintenance training program must have the capability to detect errors and perform easier repair and maintenance works for the system. | EX
1.3. Conditions of Use
1.3.1. | MUST | The containers and high-pressure tents must be weather-resistant and usable in outdoor temperatures ranging from -35 °C to +40 °C, in accordance with A1, A2, A3, B1, B2, B3, C0, C1, M1 & M2 of STANAG 4370. The tenderer shall provide the testing of the container prior to the delivery of the container. | MD | This requirement applies to the mechanical component. Indoor climate requirements are specified in sections 1.3.2–1.3.4.
1.3.2. | MUST | The containers and high-pressure tents must enable to store items at a temperature of at least +20 °C to a maximum of +25 °C throughout the year, special requirements shall be described in detail for each container separately. | EX | All HVAC systems are integrated to container system and can be used in work and in transport position. Including heating, cooling and air dehumidification
1.3.3. | MUST | When storing and transporting packed-up containers, the inside temperature and humidity must enable to preserve the equipment and appliances (incl. information technology hardware) throughout the year in a way that guarantees the immediate use of the equipment and appliances. The tenderer shall describe their solution in detail. | EX | An air dryer will be integrated into the container in accordance with the requirements. In addition, there is a heater unit to maintain the required temperature.
1.3.4. | MUST | Maximum dust resistance must be guaranteed, according to the IP54 protection class, for the container that is packed up or in the working position in accordance with the solution offered by the tenderer, which must be coordinated with the contracting authority. | EX
1.4. Requirements for Container Transport and Installation
1.4.1. | MUST | It must be possible to transport the containers using all main means of container transport. The tenderer shall describe their solution in detail. | EX | Container can be transported by road, rail and sea.
1.4.2. | MUST | It must be possible to transport and install the containers in difficult terrain (outside the road network) and weather conditions (precipitation, freezing, strong wind). The tenderer shall describe their solution in detail. | EX | Container system is designed to withstand shocks and vibrations that occur on transport on difficult terrain
1.4.3. | MUST | It must be possible to use the containers when installed directly on the ground. | EX
1.4.4. | MUST | The access routes of the two-in-one and three-in-one containers must enable to access the container in both positions, i.e. the transport and opened position. | EX
1.4.5. | MUST | It must be possible to install and open containers with one or two sides opened on the ground without additional support (according to clauses 1.1.4 and 1.1.5). | EX
1.4.6. | MUST | The containers must be minimally equipped with a manual lifting system, operated with a hand-crank and drill. The operating point for the legs (place where the crank and drill are attached) cannot rise higher than 1.5 metres. The lifting system must be certified and tested and its parameters must correspond to the mass of the container. The lifting system must enable to lift the container at least 1750 mm from the ground, each leg must be able to endure a third of the container's mass. The lifting equipment cannot exceed the overall dimensions of the container in the transport position. The tenderer shall describe their solution in detail. | MD | Containers are equipped with manual lifting system from Haacon GmbH that can lift the container up to 1750mm. Modification of the cranc mechanism to keep in lower than 1500mm is included in the project
1.4.7. | MUST | The floor of containers with one or two parts opened must be on one level. | EX
1.5. Requirements for Construction and Materials
1.5.1. |  | Containers
1.5.1.1 | MUST | Frame: must be made of S355 steel, other steel parts of S355 or S235 steel. The environmental class of the load-bearing structures and external boundaries is at least C3. | EX
1.5.1.2 | MUST | Surface treatment: the surface treatment of the container must be performed in accordance with standard ISO 12944 and scheme EPPUR160/3-FeSa2½, colour RAL8011 RAL6031 , matte. The environmental class of the surface treatment must be at least C3. | EX
1.5.1.3 | MUST | Gaskets: the gaskets used in the containers must be made of EPDM rubber with a hardness of 40-90 and must ensure maximum dust resistance. | EX
1.5.1.4 | MUST | External walls: three-layer sandwich panels with a PIR/PUR filling. The outer and inner surfaces of the sandwich panels are made of 0.5–0.6 mm steel sheets with a polyester coating (S280+Z275 PE RR20) which are glued to a foam core. Fire resistance at least REI 30. The materials used must be safe for people and the environment. The exterior surface of the containers do not need to be made of metal sheets; they can also be covered with a UV-resistant fibre glass laminate, which gives equivalent strength properties as a metal sheet. | EX
1.5.1.5 | MUST | Floors: structurally similar to the walls and ceiling, the estimated use load is at least 3.0 kN/m². Floor strength in all containers 43 Heavy according to classification ISO:10874; resistance R1=<10⁸Ω, R2=<10⁸Ω according to classification EN 1081; antistatic =<2kV in accordance with classification EN 1815. Fire resistance at least REI 30. The flooring is PVC with welded joints. The sub-floor exterior surface of the container can be made of UV-resistant fibre glass laminate, which gives equivalent strenght properties. | EX
1.5.1.6 | MUST | Interior rooms must have floor heating of at least 1 kW. Technical rooms, which store water, must have floor heating of at least 200 W. Floor heating must be equipped with a thermostat. Floor heating must be covered with a material that distributes heat inside the room equally. Floor heating shall be implemented using heating cables rated at 25 W/m, with the total installed cable length sized to meet the required minimum heating capacities stated above. The system shall be capable of maintaining a minimum floor surface temperature of +10 °C at the central floor area, including in transport mode. | EX
1.5.1.7 | MUST | Roof/ceiling: structurally similar to the walls and floor, the estimated snow load of the roof is at least 2.0 kN/m². Roofing must have UV protection and be able to bear walking on the roof and the removal of snow. Walking on the side roofs of an opened container (two-in-one and three-in one) is not intended. For ceilings and walls, it must be possible to attach accessories with magnets. For each container, at least 10 magnetic hooks must be delivered, which has a carrying capacity of at least 5 kg when attached to the wall or ceiling. | EX
1.5.1.8 | MUST | Walls, floor and ceiling with a thermal conductivity of λ < (below) 0,028 W/mK (EN 13165). | EX
1.5.1.9 | MUST | Exterior door: clear opening min 1015 × 2035 mm 1100 × 2100 mm metal profile door (EI 15) with a door closer, mechanical lock and handle (exact type to be coordinated with the contracting authority). Door sills RST. U = 1.4 W/m2K. | EX
1.5.1.10 | MUST | Insulation: must ensure a container's heat loss of no greater than 3.5 kW. | EX
1.5.1.11 | MUST | Minimum thicknesses of panel insulation (mm): | EX
1.5.1.11.1 | MUST | Two-in-one and three-in one container | EX
|  | Main container mm End walls 50 Roof 60 Floor 90 Walls 50 Niche covers 50 Utility room partition wall 50 Opened sections Floors 70 Roofs 50 Side walls 50 End walls 50 | EX
1.5.1.11.2 | MUST | Laboratory containers and connection containers: | EX
|  | Main container mm End walls 60 Roof 60 Floor 75 Walls 60 | EX
1.5.1.11.3 | MUST | Med. logistics container with hooklift system: | EX
|  | Main container mm End walls 100 Roof 100 Floor 100 Walls 100 | EX
1.5.1.11.4 | MUST | Resource container 1 and 2 | EX
|  | For resource containers, oxygen- and water-producing rooms must have walls with the following specifications: PIR/PUR foam and λ < (less than) 0,028 W/mK (EN 13165); generator room and dividing wall between the generator room and production room must have an acoustic rock wool panel, 100 mm: λ < (less than) 0,038 W/mKe. | EX
1.5.2. |  | Platforms | EX
1.5.2.1. | MUST | The platforms will be connecting together containers and tents, with measurements 1500 x 2200 mm 1500x2090 mm, made from aluminium. The laboratory container, connection container and ward container will have platforms. Other containers have doors. | EX
1.5.3. |  | High-pressure tents | EX
1.5.3.1. | MUST | Max pressure in the high-pressure tents cannot exceed 8 bar. The high-pressure tents must include removable and washable fabric used for insulation. It must be possible to connect the tents with doors. If tents are not connectable directly, they must be connectable with a corridor stipulated in clause 1.5.4. In such an event, there must be 4 pcs of additional corridors in addition to what's listed in clause 1.1.6. | EX
1.5.4. |  | High-pressure corridors | EX
1.5.4.1. | MUST | Max. pressure in the high-pressure corridor cannot exceed 8 bar. The high-pressure corridor must include removable and washable fabric used for insulation. It must be possible to be easily and quickly (the tenderer shall describe their solution in detail) connect and disconnect the corridor from the container or tent. The corridor must provide a cover for the passage from the sides and underneath and be weather-proof. The length of the corridor must be in compliance with the length of the passage. Corridors must be universal for both the container-to-container connection and tent-to-container connection. | EX | Containers have special connection frame where the corridor is fastened with zippers and velcro flaps
1.6. Utility Systems
1.6.1 |  | Low voltage systems (electrical system of consumers (excl. resource containers)) | EX
1.6.1.1. | MUST | The electrical system of each container must be designed based on similar principles to the greatest extent possible (same input supply, wiring, lighting etc.) if the requirements of the container do not require a more specific design. | EX
1.6.1.2. | MUST | The electrical systems must be designed and constructed according to the Estonian standard series EVS-HD 60364 or Lithuanian standard series LST HD 60634 (as per the buyer). | EX
1.6.1.3. | MUST | The input of each container must be in TN-S wiring system (further internal separations in the container either IT or other wiring system as per requirements). This will be specified during the design phase. The power supply of the MMH must work with either generators or the national power grid. Switching must have RLA switches. | EX
1.6.1.4. | MUST | Each container kit must include sufficient electrical equipment (cables with the required length, switchgear, power sockets etc.) to achieve the electrical connections between containers and with the resource containers, while ensuring maximum safety of the equipment and personnel. Appropriate protection methods must be used (grounding, equipotential bonding, circuit breakers, residual current breakers etc.). The tenderer shall describe their solution in detail. | EX | All containers are equipped with correct size extension cables and grounding kits to assure grounding, equipotential bonding between container and also power supply from ressource modules
1.6.1.5. | MUST | Each container must allow to continue the electrical supply to the next container. H07-type rubber cable must be used for outdoor connections. Each container must have a sufficient amount of inputs and outputs to achieve the general power supply system of the MMH. | EX
1.6.1.6. | MUST | In addition to the inputs and outputs of the general electrical circuit. Each container must have at least four F-Type 16A IP67 and two at least one 16A 32A 400V outputs (if not specified otherwise). The IP rating of outdoor devices is at least IP65. Indoor IP ratings must be designed according to the Estonian standard series EVS-HD 60364 or Lithuanian standard series LST HD 60634 and in coordination with the contracting authority. | EX
1.6.1.7. | MUST | The containers are categorized into medical group classifications in accordance with the Estonian standard EVS-HD 60364-7-710 or Lithuanian standard LST HD 60364-7-710 (as per the buyer) and must comply with the set requirements of the corresponding medical group. | EX
1.6.1.8. | MUST | The electrical system of each container must include a phase monitoring and sequence test device and an electricity meter. | EX
1.6.1.9. | MUST | The MMH must be provided with separate additional equipment: cable protection ramps, extension cord, electrician’s tool set. 16 pcs of cable protector with a width of 600 mm, height 70 mm, length 1000 mm, 3 channels, capacity of at least 5 tonnes, black or dark green; 8 pcs of hose bridge with a width of 500 mm, height 90 mm, length 850 mm, 2 channels 75x70 mm. | EX
1.6.1.10. | MUST | The workspaces of the container must be provided with general lighting in accordance with legislation (refer to Estonian standards EVS-EN 12464-1/ Lithuanians standards LST EN 12464-1 and EVS-EN 12464-2/ LST EN 12464-2). General lighting must be DALI impulse dimmable. The utility room must also be illuminated. | EX
1.6.1.11. | MUST | The lighting system must enable switching to emergency and battle lighting mode in modules agreed with the end user. Lighting must be manually regulated between three regimes – regular lighting, combat lighting and off. Combat lighting means the application of blue light when an exit is opened, regular lighting is restored when closed. | EX
1.6.1.12. | MUST | In order for the electrical system of the containers to function, the discharge of the batteries during 6 months of storage without power supply must be excluded. At the same time, the solution should ensure that when external power is connected, the batteries are charged without any additional switching. | EX
1.6.1.13. | MUST | All electrical installations must be inspected and tested thoroughly according to local legislation. Compliance audits must be conducted with the declarations of conformity for the electrical installation. | EX
1.6.1.14. | MUST | All cables, switches, cabinets, power sockets, plugs, generators and all other electrical devices must be marked with legible and durable clips, stickers or other marking materials. | EX
1.6.1.15. | MUST | All containers and parts of the MMH must have up-to-date electrical documentation available in in corresponding cabinets. | EX
1.6.2. |  | Extra low voltage & data communication systems
1.6.2.1 | MUST | All inputs and outputs of the connection board: wire communications terminals; MIL-DTL-38999 RJ-45 inputs and outputs. | EX
1.6.2.2 | MUST | All exterior LAN pins (RJF sockets) must be equipped with gas-powered overvoltage limiters, i.e. container sockets which are designed for connecting with exterior devices or other containers. | EX
1.6.3. |  | Heating and cooling | EX
1.6.3.1 | MUST | The solution must be equipped with diesel heaters and coolers (pre-heatable/pre-coolable) which must operate on 24V DC supply voltage that ensures a quick temperature rising or lowering to 20 °C within 15 minutes from setting the container into the working position (except logistics container). This requirement does not apply, if the system has no possibility of checking the containers' interior temperature. The tenderer shall describe their solution in detail. | EX | Containers are equipped with Webasto AirTop 55D diesel heater to assure quick temperature rise and additional heating capacity when full lectrical power is not available. Diesel heater cirulates internal air in order to acheive faster heating on the container interior.
1.6.4. |  | Water supply and sewerage | EX
1.6.4.1. | MUST | Designed and constructed according to the interior and purpose of the container. The container must be able to pump in water from an external water source. | EX
1.6.4.2. | MUST | Containers that have water connections must be delivered with sewerage and clean water hoses with a length of at least 25 m and thickness of at least 1 inch. The hoses' technical parameters must be in accordance with volumes that are planned for hoses by the containers. Sewerage hoses must be with claw couplings and clean water hoses with Storz couplings. Hoses must have heating cables and heating insulation must be delivered for them. The sterilization container's sewerage hose can be without heating. | EX
1.6.5. |  | Waste management | EX
1.6.5.1. | MUST | The tenderer must have a waste solution for each container separately and the whole system, while taking into account domestic and medicinal waste and sharp objects separately. | EX | a) Medical waste (infectious waste, sharp objects, medicines, and all materials contaminated with blood or body fluids, which must be separated from other waste to avoid cross-contamination); b) Non-medical waste (includes general waste such as packaging materials, food waste, non-contaminated office supplies, etc.); Waste is collected in color-coded bags/containers according to the purpose of the module and the content of the medical procedures performed there: Red bags/containers for infectious waste; Yellow containers for sharp objects such as needles and syringes (special puncture-proof containers are designed to dispose of needles, scalpels, and other sharp objects); Black or transparent bags for general waste. In the IMM and WDM modules, the trash bag holders are integrated into the furniture. In addition, the necessary number of trash carts with lids, separately marked trash cans, and needle containers are installed in the modules, respectively, and are clearly marked (https://www.gimaitaly.com/DocumentiGIMA/PaginePDF/EN/74_ENG.pdf). The waste is stored in accordance with the designated collection area agreed upon during the design phase and the requirements for proper subsequent waste management.
1.6.6. |  | Illumination | EX
1.6.6.1. | MUST | Each container must have a combat lighting regime that can be switched off (i.e. application of blue light when opening an exit). Each container must have a main switch near the entrance. | EX
1.6.7. |  | Monitoring system | EX
Automatic parameter monitoring system – required minimum monitoring in each container: temperature and air humidity inside and outside the container, the temperature of refrigerator in two points if the refrigerator is in the container. The system must be modular in order to add other parameters, if possible. The kit must include a 27” IP44 touch monitor that can be cleaned and disinfected. The system must be able to send notifications to e-mail if: the temperature of the container is low, the battery voltage is low, a power outage occurs. Storage areas for medicinal products must be traceable according to GDP Guidelines. The system must be equipped with a Cat6 data cable with the endpoint in a weak current cabinet. Add measurable electricity consumption, current, voltage, and battery system voltage. Each container has at least two temperature sensors and one humidity sensor. The sensors are installed in the container extremes or a container GDP-based temperature mapping is performed to select the most suitable location for the sensor, based on the container's nature. The system measures external temperatures with at least two sensors from different sides of the container and humidity with at least one sensor. The external temperature must be presented as the arithmetic average of two measurements. The system must measure the temperature of each container's refrigerators with at least one sensor and at least two sensors in the pharmacy container (except for the refrigerator where only one extreme is defined, based on GDP). A graphical user interface must be available for system management, value reading, historical record viewing (graphically), configuration and limit setting, accessible over the container network from a secure network address via internet browsers. The user interface must be easy to read and understand. The user interface must be used to define alarms, their limits, view measurement history, and export them as a CSV file to a memory stick. The value history is kept in CSV format for at least two years. The user interface must be able to assign users with different rights and roles (at a minimum, data reading and configuration). Users must be able to set passwords (21CFR). The system must be easily expandable and must be able to add containers, such as measuring the generator's fuel level and water tank level. The automation system must be installable on a DIN rail in an electrical panel, and the automation system's power must be provided by the container's autonomous power supply for at least 24 hours. The automation system must have an internet connection, through which the system can send alarms to email addresses that can be specified from the user interface. The user must be able to define the alarms to be sent from the measurable measurement points measured by the system by specifying the alarm criteria for them.
1.6.8. |  | Main Systems | EX
1.6.8.1. | MUST | The main systems are systems whose non-functioning is considered to not impair the functionality of the container for several hours. For example, main systems do not include boilers or similar equipment. | EX
1.7. Normative Documents
1.7.1. | MUST | All the parts and components of the container must comply with the legislation, regulatory documents, rules and standards applicable in the Republic of Estonia/Lithuania and the European Union. | EX
1.7.2. | MUST | The sides and top of the container must bear medical markings in accordance with the Geneva Convention in a way that enables to easily cover and uncover the markings. | EX
1.7.3. | MUST | Working environment requirements (noise, vibration, lighting, ergonomics, ventilation, work safety) must comply with the norms applicable in the Republic of Estonia/Lithuania and the European Union. | EX
1.7.4. | MUST | The frame of the container must be designed and constructed in accordance with the following ISO standards: | EX
1.7.4.1 | MUST | ISO 668; Series 1 freight containers. Classification, dimensions and ratings. | EX
1.7.4.2 | MUST | ISO 1161; Series 1 freight containers. Corner and intermediate fittings. Specifications. | EX
1.7.4.3 | MUST | ISO 3874; Series 1 freight containers. Handling and securing. | EX
1.7.4.4 | MUST | ISO 1496-1:2014; Series 1 freight containers. Specification and testing. | EX
EX
2. INTERMEDIATE MODULE, LABAROTORY, DIAGNOSTICS AND TRIAGE SET
2.1. Contents of a Set
2.1.1. | MUST | Laboratory container | EX
2.1.2. | MUST | Connection container | EX
2.1.3. | MUST | High-pressure tent | EX
2.1.4. | MUST | High-pressure corridors for container-to-container or container-to-tent connections | EX
2.2. General
2.2.1. | MUST | The set must include all the equipment required for handling the containers, high-pressure tent and high-pressure corridors. | EX
2.2.2. | MUST | It must be possible to connect the container with other containers of the same type to form a single system. | EX
2.2.3. | MUST | The laboratory container and triage tent system must consist of an ISO 1CC 20' container and high-pressure tent. The container and tent must have a shared power supply and communication system. | EX
2.2.4. | MUST | The containers shall have four entrance platforms to which the tent or a container can be attached (see clause 1.5.2). | EX
2.3. Low Voltage Systems
2.3.1. | MUST | The electricity mains of the container is 3-phase 400V 50Hz. The container must be powered from an external supply (grid or generator). In case there is no external supply for the generator the main systems (including lighting) of the container, it shall be powered by batteries, whereby the batteries shall provide at least an hour at full power. | EX
2.3.2. | MUST | In addition to the 400V and 230V electrical systems, there must be 12V and 24V electrical systems, if this is required for the equipment used. The container must be equipped with the required low-voltage/extra low-voltage transformers. | EX
2.3.3. | MUST | The electrical system of the container must be designed and constructed to provide 3-phase 32A 3A 400V outlets for four medical containers. | EX
2.3.4. | MUST | The interior power system of the container must be designed to meet the needs of the planned equipment (power outlets types and quantity, IP ratings, lighting, potential equalization sockets, circuit breakers, protection devices etc.) according to all relevant standards and be appropriate for the purpose of the container. The tenderer shall describe their solution in detail. | MD | Electical system design will be verified after container design is finalized. Considering power consumption, protection level needed for the equipment etc.
2.4. Extra Low Voltage (ELV) & Data Communication Systems
2.4.1. | MUST | The container must contain RJ45 outlets (eight in total) at workstations and at least eight two inputs with RJF type sockets in the external feedthrough cabinet, with a possibility to close them with a cap or a stopper. The cables have an endpoint in the ELV & data communication cabinet on the RJ45 panel and they are marked. Cat6 twisted pair cable is used, the colour code of which is T568B. Cables are marked at both ends. | EX
2.4.2. | MUST | Workstations (8 pcs) shall be provided with fibre optic cable (standard OM-3 or OM-4) with two fibres and terminating in LC duplex adapters; in the extra low-voltage cabinet, fibres shall be run to the panel of FO cables, with endpoints in LC adapters and marked at both ends. | EX
2.4.3. | MUST | An SM fibre optic fibre pair (standard OS-2), with eight pairs of fibres in total, shall be run to the external feed panel; the fibres shall be terminated with LC duplex adapters on the communication cabinet panel. On the outer feed panel, a 4ch Stratos HxMA type socket is used; it is closed with a stopper. Plug/socket standard MIL83526/MIL20/MIL21. | EX
2.4.4. | MUST | The ELV & data communication cabinet must be at least 19-inch, which must contain at least an 18-port Cisco switch (or equivalent, programmable according to the same principles) for setting up tunnels, and a Cisco 4G router (or equivalent, programmable according to the same principles) with 4G and a GPS antennas on the roof. the ELV & data communication cabinet must be provided with proper cooling. | EX
2.4.5. | MUST | the ELV & data communication cabinet must have a server computer to run the monitoring system, generator and medical equipment software. The cabinet must be provided with proper cooling. | EX
2.5. Illumination
2.5.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. The tenderer shall describe their solution in detail. | EX | LED lights for exterior are integrated to container systems
2.5.2. | MUST | Interior lighting shall provide uniform and diffused lighting throughout the room. Lighting must enable dimming and DALI control. Lighting on working surfaces at 950 mm from the floor surface shall not be less than 1000 lx. | EX
2.5.3. | MUST | The illumination of technical areas and rooms must be at least 200 lx if not stated otherwise. | EX
2.5.4. | MUST | The interior and external lightning of the container shall be ensured by the generator, external power or batteries, whereby the batteries shall provide at least an hour at full power. | EX
2.6. Ventilation
2.6.1. | MUST | The system must ensure proper air exchange, air intake with a heat exchanger of at least 30 l/s. Ventilation systems are described in Role 2 HVAC project. | EX
2.6.2. | MUST | The container must include integrated stainless steel furniture, the exact solution of which will be coordinated with the contracting authority during the design phase. The furniture shall include at least 2000 L of storage space, including four drawers with a capacity of approx. 14 L, 15 single-door lockers with a capacity of 16 L, four double-door lockers with a capacity of 24 L, two 80L cabinets for waste bins. | EX
2.7. Water and Sewerage System
2.7.1. | MUST | The container shall contain a two-tap hand-washing station, measuring at least 580 mm in width, not more then 200 mm in length and at least 100 mm in depth. Stainless steel sink for hand washing with rounded corners. A container for clean water with a capacity of at least 50 L is required. The system must have a boiler with a capacity of at least 10 L and a power of 2 kW. A pump must be installed at the hand washing station to draw clean water into the container and create water pressure in the system. A water pipe must be brought to the outside of the container, through which it can be permanently or temporarily connected to an external water source for pumping clean water into the container. Used water will be discharged through a pipe on the outside of the container, which can be connected temporarily or permanently to a sewage or used water container. Connecting to a clean water source and a sewage system must be possible independently of the season. | EX
2.7.2. | MUST | The handwashing station must be integrated with the furniture installed in the container. | EX
2.8. Medical Gases
2.8.1. | MUST | The container must be equipped with a medical gas distribution system and a possibility for external O2 connection to create a network between containers (surgical, intensive care and emergency medical containers, triage tent). Oxygen outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. For gas network connection, the container must be provided with an external oxygen inlet connection and be supplied with an oxygen hose of at least 25 m in length and an inside diameter of 8 mm. | EX
2.8.2. | MUST | The container must be equipped with a medical gas distribution system and a possibility for external pressurised medical air connection to create a network between containers (surgical, intensive care, emergency medical, treatment containers, triage tent). Pressurised outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. For gas network connection, the container must be provided with an external pressurised medical air inlet connection and be supplied with an oxygen hose of at least 25 m in length and an inside diameter of 8 mm. | EX
2.8.3. | MUST | Both sides of the container must have a gas panel installed, each of which has 4 SS-type or 4 DIN- type oxygen quick connect sockets and 2 SS-type or 2 DIN- type medical compressed air quick connect sockets. The panels must be interconnected and the connection must allow for the transfer of gases from one panel to the other without obstruction. | EX
2.9. Equipment to Be Included in the Laboratory, But Not to Connecting Containers
2.9.1. | MUST | Warming cabinet for infusion solutions – with a capacity of at least 20 L. Internal temperature 39 °C +/- 2 °C. | EX
2.9.2. | MUST | Refrigerator for medicines and solutions – with a capacity of at least 40 L, intended for storing medical preparations and medicines. It must be possible to set the temperature with an accuracy of 1 degree in the range from +2 to +8 °C. The refrigerator must be divided into two sections with one shelf, and the shelf must withstand a load of at least 30 kg. The noise level of the appliance must be less than 50 dB(A) and the weight must be less than 60 kg. | EX
2.9.3. | MUST | Transfer board – weighing no more than 4 kg. Dimensions at least 1700 x 600 mm, made of high-density polyethylene, easy to clean and resistant to acid cleaning agents, must not cause any artefacts on X-ray images. | EX
2.9.4. | MUST | Two Four plasma thawing devices that can each thaw and warm up to four bags of plastic plasma at a time, and also have a mode for emergency thawing and warming of up to eight bags of plasma each. Capacity of at least 9 liters. Temperature range of at least 37–45 °C. Device weight not exceeding 18 26 kg. Separate temperature programs for blood and plasma products. Audio-visual alarm that notifies the end of the program. Dimensions no larger than 340x600x320 470 x 413 x 375 mm (must fit and be operable on the work surface of an connection container). | EX
2.10. Laboratory Container with Triage Tent, Additional Requirements
2.10.1. High Pressure Tent
2.10.1.1. | MUST | The floor area of the triage tent in the working position must be at least 60 m². | EX
2.10.1.2. | MUST | The high-pressure tent must have two doors, one at each end, and two side doors, the dimensions and positions of which will be coordinated in the design phase with the contracting authority. Ventilation systems are described in Role 2 HVAC project. | EX
2.10.2. Container Equipment
2.10.2.1. | MUST | In vitro diagnostics medical devices must comply with Regulation (EU) 2017/746 of the European Parliament and of the Council of 5 April 2017 on in vitro diagnostic medical devices and repealing Directive 98/79/EC and Commission Decision 2010/227/EU. | EX
2.10.2.2. | MUST | Portable rapid haematology analyser – supporting the following rapid tests: sodium (Na), potassium (K), chloride (Cl), anionic balance, ionic calcium (iCa), glucose (Glu), urea nitrogen (BUN)/urea, creatinine (Crea), lactate, haematocrit (Hct), haemoglobin* (Hgb), pH, PCO2, PO2, TCO2, HCO3, base excess (BE), sO2, ACT Kaolin, ACT Celite, PT/INR, β-hCG, cTnI, CK-MB, BNP. | EX
2.10.2.3. | MUST | Chemistry analyser – supporting the following rapid tests (results in less than 15 minutes): ALB, ALP, ALT, AST, BUN, Ca, CI-, CRE, GLU, K+, Na+, TBIL, tCO2, TP, CHOL, CHOL/HDL*, HDL, LDL, TRIG, VLDL, ALT, AST, GLU, AMY, GGT, TBIL, ALB, ALP, ALT, AMY, AST, BUN, Ca, CRE, GGT, GLU, TBIL, TP, UA, ALT, AST, BUN, CRE, GGT, GLU, Cl-, K+, Na+, tCO2, BUN, CRE, BUN, CK, Cl-, CRE, GLU, K+, Na+, tCO2, ALB, BUN, Ca, Cl-, CRE, GLU, K+, Na+, PHOS, tCO2, BUN, Ca, Cl-, CRE, GLU, K+, Na+, tCO2, Mg, lactate dehydrogenase, DBIL, TBIL, TP. | EX
3. OPERATION MODULE (Surgery Container)
3.1. General
3.1.1. | MUST | The walls and floor of the room should be easy to clean and resistant to disinfectant solutions. The resistance of the floor material must be in the range of 104–106 Ohms. The tenderer shall describe their solution in detail. | EX | Walls interior is painted steel to easily clean it. Flooring is antistatic PVC glued covering and joints are welded
3.1.2. | MUST | The container must include an emergency exit for leaving the working space of the container when the front door is obstructed; location shall be specified during the design stage. | EX
3.1.3. | MUST | The container must provide automated disinfection (disinfection of all surfaces in less than 15 minutes), the device must be contained inside the container. The tenderer shall submit their own solution. | EX | A container is equipped with a UV-C air purification device NEXA PROLUX G 30WA/SPH02, and an additional autonomous and fully automatic atomizer the Anios Aerosept Ultra 15 device intended for the disinfection of previously cleaned surfaces.
3.2. Low Voltage Systems
3.2.1. | MUST | The electricity mains of the container is 3-phase 400V 50Hz. The container must be electrically autonomous, which is ensured, in addition to the possibility of connecting to the external power supply (i.e grid), by a diesel generator with a capacity of at least 8 kW. The fuel supply of the diesel generator must ensure the autonomy of the container for 36 hours. The fuel tank of the container must allow refuelling from outside the container. | EX
3.2.2. | MUST | During the operation of the container's electric generator, it must be possible switch to an external power supply and back without a power interruption. In the case of power interruption, the autonomous electric generator must start and connect to the network within 30 seconds. In case there is no grid connection and the operation of the generator is excluded the main systems (including lighting) of the container shall be powered by batteries, whereby the batteries shall provide at least an hour at full power. Switching with an RLA switch. | EX
3.2.3. | MUST | The standalone generator must start automatically. At startup time, uninterrupted operation of the system must be ensured (e.g. UPS). The container must have a built-in fuel tank for 36 hours of medium-load work, refueling must be done from outside the container. | EX
3.2.4. | MUST | It must be possible to use the power of the autonomous electric generator to power additional containers through output connectors. | EX
3.2.5. | MUST | In addition to the 400V and 230V electrical systems, there must be 12V and 24V electrical systems if this is required for the equipment used. The container must be equipped with the required low-voltage/extra low-voltage transformers. | EX
3.2.6. | MUST | The interior power system of the container must be designed to meet the needs of the planned equipment (power outlets types and quantity, IP ratings, lighting, potential equalisation sockets, circuit breakers, protection devices etc) according to all relevant standards and be appropriate for the purpose of the container. The classification of the container in accordance with Estonian standard EVS-HD 60364-7-710 or Lithuanian standard LST HD 60364-7-710 is group 2 (as per the buyer). | EX
3.3. Extra Low Voltage & Data Communication Systems
3.3.1. | MUST | The container must contain ELV outlets (14 in total) at workstations and at least two ELV inputs with RJF type sockets in the external feedthrough cabinet, with a possibility to close them with a cap or a stopper. The cables have an endpoint in the ELV & data communication cabinet on the RJ45 panel and they are marked. Cat6 twisted pair cable is used, the colour code of which is T568B. Cables are marked at both ends. | EX
3.3.2. | MUST | The ELV & data communication cabinet must be at least 10-inch, which must contain at least a 16-port Cisco switch for setting up tunnels, and a Cisco (or equivalent, programmable according to the same principles) 4G router with 4G and a GPS antennas on the roof. The cabinet must have a server computer to run the monitoring system, generator and medical equipment software. The cabinet must be provided with proper cooling. | EX
3.4. Illumination
3.4.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | MD
3.4.2. | MUST | Interior lighting must ensure uniform and diffused illumination, meeting the requirements for surgical premises, throughout the room. Illumination must support dimming and DALI control. The illumination of working surfaces at 950 mm from the floor must be at least 1000 lx. | EX
3.4.3. | MUST | The illumination of technical areas and rooms must be at least 200 lx if not stated otherwise. | MD
3.4.4. | MUST | The interior and external lightning of the container shall be ensured by the generator, external power or batteries, whereby the batteries shall provide at least an hour at full power. | EX
3.5. Ventilation
3.5.1. | MUST | The container must have proper air exchange and it must be equipped with a HEPA 13 filter. Two-stage air purifier, the first stage at least G4 and the second at least F9. Mechanical intake and electric heating with a calorifier. Indoor air in winter: 18–25 °C (+/- 1 °C), indoor air in summer: 24 °C. | EX
3.5.2. | MUST | The air in the container must comply with ISO 7 cleanroom requirements, with UV power of 110 W, air purification rate of 140 m3/h. | EX
3.5.3. | MUST | The room must have an overpressure of up to 20 Pa, air humidity in the range of 4–60% (reachable from 100% humidity in less than 30 minutes), air exchange at least 25 times per hour, including at least 10 times with ambient air. It must be possible to clean the volume of the container with a UV air purifier every 60 minutes. Air temperature must be selectable in the range of +18...+25 °C with a tolerance of +/- 1 degrees at an outside temperature range of -35 to +40 °C, assuming that the room includes up to eight people and the specified equipment. | EX
3.6. Water and Sewerage System – None
3.7. Oxygen System
3.7.1. | MUST | The container must include gas pipes and a possibility of an external O2 connection to establish a network between containers. Oxygen outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. For connection to the gas network, the container must include an oxygen hose of at least 10 m in length with an internal diameter of 8 mm. | EX
3.7.2. | MUST | The container must be equipped with a medical gas distribution system and a possibility for external pressurised medical air connection to create a network between containers (surgical, intensive care and emergency medical containers, triage tent). Pressurised outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. For gas network connection, the container must be provided with an external pressurised medical air inlet connection and be supplied with an oxygen hose of at least 25 m in length and an inside diameter of 8 mm. | MD
3.7.3. | MUST | On both sides of the container, a gas panel must be installed, each with 4 SS 8752430 or 4 DIN 13260-2 type oxygen quick connect sockets and 2 SS 8752430 or 2 DIN 13260-2 type medical compressed air quick connect sockets. The panels must be interconnected and the connection must allow for the transfer of gases from one panel to the other without obstruction. The dimensions of the panel must be chosen according to the quick connections used, but must not be larger than 800x600x300 mm. | MD
3.8. Equipment
3.8.1. | MUST | All medical devices must comply with Regulation (EU) 2017/745 of European parliament and of the council of 5 April 2017 on medical devices. | EX
3.8.2. | MUST | Aspirators – two devices: air flow rate at least 45 L/min., maximum vacuum at sea level -91 90 kPa- 900 mbar- 675 mmHg under normal conditions, vacuum level accuracy: -1–0 bar ± 16 mbar (Class 1.6), possibility to use 2 x 3L or 2 x 5L 2L/4L/5L transparent secretion tanks, possibility to use ø 6mm 8mm or ø 10mm hoses, maximum power not more than 200 W, weight with secretion tank not more than 24 kg, noise level < 48 dB (A) @ 1 m (ISO 7779). | EX
3.8.3. | MUST | Oxygen generator – with a capacity of at least 20 L/oxygen per minute, with a pressure of 5 bars and an oxygen concentration of 93%. The weight of the oxygen generator must not exceed 120 kg. The current consumption shall not exceed 1.4 kVA. The oxygen generator must be single-phase. The oxygen generator must comply with the IP rating of 42. The oxygen generator must have three oxygen outlets, including at least one with adjustable pressure between 0 and 8 bars; one outlet must support flow rate adjustment and one outlet must be designed to connect the device to an oxygen bottling system. The current oxygen pressure must be indicated by a manometer. The generator must have a separate outlet for discharging oxygen-depleted air. It must be possible to remove the oxygen generator from the container if necessary. The oxygen generator must be supplied with a packing box or boxes with MIL-15 certification and IP65 protection rating. The oxygen generator must be provided with a connecting hose for connecting the generator to the DIN 13260-2 or SS 8752430 standard quick coupling. A mounting point for a 10L exchangeable oxygen cylinder and connections to the compressor, such that the oxygen supply from the cylinder continues automatically in case the compressor stops working. | EX
3.8.4. | MUST | Medical air compressor – with an output capacity of at least 40 LPM, output pressure not exceeding 7 bar, noise level not exceeding 53 dBA, at least Class 0 (ISO 8573-1:2010) air purity. Continuous air production at least 100 LPM, configurable from 0 to 10 bar. Connected to pipework described in clause 3.7.2. | EX
3.8.5. | MUST | Operating table – two operating tables one operating table, must be integrated in the container. The operating table shall be suitable for advanced mobile field structures and transportable in dedicated transport cases by military or civil defence vehicles, aircraft, helicopters, rescue vehicles, or equivalent transport platforms. The length of the operating table should be at least 2100 x 560 mm 2000 x 500 mm or equivalent, with a weight capacity of at least 320 250 kg dynamic load or equivalent. The height of the table must be adjustable within a range of at least 680–1120 mm, with lateral tilt in both directions of -35° to +75° and a minimum of 40° Trendelenburg and Antitrendelenburg position. The backrest should be adjustable within a range of -35° to +75°, the legrest within a range of -90° to +20°, and be able to rotate 180°. The height of the table must be adjustable for surgical procedures in field conditions, mechanically, hydraulically, electrically, or by equivalent assisted system. The backrest should be adjustable within a range of -35° to +75° -45° to +75°, the legrest within a range of -90° to +20°, and be able to rotate 180°.The table must have a resuscitation function, meaning the table top can be quickly lowered to a horizontal position. The table should be composed of at least four parts (detachable headrest, backrest, and legrest). The table should be composed of at least five parts, including head section, backrest, seat section, and split leg sections. The table top must be compatible for use with X-rays. All components of the operating table must have accessory rails. Side rails approximately 25 x 10 mm or equivalent shall run along both sides of the table and be suitable for mounting surgical accessories. All components of the operating table must have accessory rails. The operating table must come with trauma surgery equipment. The frame of the operating table must be made of stainless steel. The frame, base, column, and trolley of the operating table must be made of stainless steel or equivalent corrosion-resistant material suitable for field medical use. The table must be disconnected from the power supply during the procedure. The lifting, lowering, and positioning of the table must be done electrically, and the table must have a built-in battery.The lifting, lowering, and positioning of the table may be mechanical, hydraulic, gas piston assisted, electrical, or equivalent. The table must be grounded or double insulated. At least a 5m potential equalization cable must be supplied with the table. The operating table shall be supplied with a mobile transport trolley equipped with antistatic castors and braking system suitable for mobile medical operations. | EX
3.8.6. | MUST | Diathermocoagulator trolley – two devices with a power of 550 W; with nominal frequencies of 350–500 kHz, with a maximum cutting power of at least 400 200 W, maximum coagulation power of at least 200 140 W, with PCS and NE indicators, alarm and activation logs, weighing less than 10.3 kg. | EX
3.8.7. | MUST | Infusion pump – eight infusion pumps with dimensions not exceeding 245 x 90 x 180 mm, easy to use, mobile and easy to clean, weighing not more than 2.0 kg with an integrated fixation clip for the drip stand and a battery. Mains power 230 V, and the device must operate from both the mains and the lithium battery, with a battery operating time of at least 10 h 6 h at an infusion rate of 5 ml/h. The device must allow for the use of different syringes (different types of sizes 5, 10, 20, 30, 50/60 ml, recognised automatically by the infusion pump) and functions: set volume of infusion; specified infusion duration; administration of a bolus; the infusion pump must have an anti-bolus function. The device must have a coloured touch screen, indicating at least the following parameters: infusion rate (ml/h); the amount of infused liquid (at least from 0.1 to 999.9 ml); the time until the infusion is complete/the syringe emptied; occlusion level (settable within the range of approximately 300 to 900 mbar or equivalent from 0 to 1000 mbar and easily modifiable); event trend, at least 1500 per an event. The parameters must be changeable at least in the following range: rate from 0.01 to 1500 ml/h (for 10–50ml syringes); set volume of infusion from 0.1 to 100 ml; infusion duration from 1 min. to 24 hours; a keyboard shortcut must be provided to fill a bolus. The accuracy of the infusion must be at least ± 2% in accordance with the IEC/EN 60601-2-24 standards. The alarms must have both two lights representing two levels and a sound announcing at least the end of the infusion, reaching the set volume of infusion, remaining time until the end of the infusion (a few minutes prior), battery discharge, battery capacity, 230V external mains failure, technical error. The device must be equipped with audible and visual alarms indicating at least infusion completion, reaching the set infusion volume, low battery level, external mains failure, occlusion, and technical errors. Possibility to switch the syringe pumps to the alarm system (nurse call). The instructional text transmitted on the pump screen must be in Estonian or Lithuanian or English. Infrared port, in addition to data transfer, also for inserting new software upgrades. The device must be equipped with an infrared port, USB port, or equivalent interface for data transfer and software upgrades. The syringe plunger must not protrude beyond the syringe pump housing in the case of a filled syringe. Possibility to connect syringe pumps to a special quick-coupled storage compartment with power supply. | EX
3.8.8. | MUST | Monitor-defibrillator – two defibrillators with monitor, recording the physiological parameters of a patient, enabling one to monitor the following parameters of a patient: registration of a 12-lead ECG; registration of a 3- or 5-lead ECG; heart rate measurement with a measuring range of at least from 30 to 300 BPM; pulse oximetry (SpO2) with a measuring range of at least from 1 to 100%, to an accuracy of ± 2%; heart rate with a measuring range of at least from 25 to 240 BPM; monitoring of non-invasive blood pressure (NIBP) with a measuring range of at least from 40 to 270 mmHg; monitoring of invasive blood pressure (IBP) with a measuring range of at least from -30 to 300 mmHg; capnography monitoring using sidestream capnography equipped with Microstream accessories or equivalent technology the Microstream method; the device must provide feedback on the quality of performed chest compressions during resuscitation according to ECR/AHA guidelines; the device must enable round-the-clock registration of administered drugs and trend monitoring of measured parameters on the screen, as well as the printout of trends. The device must be complete with the following tools: a set of cables and electrodes for recording a 12-lead ECG; cable and sensor for monitoring pulse oximetry; different sized cuffs and hoses for non-invasive blood pressure monitoring; invasive blood pressure monitoring kit; capnography sensor; a wall mount; charger; with a protective case and closable accessory pockets (cables, electrodes, sensors, etc.) and a shoulder strap. The device screen must meet the following requirements: a diagonal screen of at least 16 cm, coloured; allow simultaneous monitoring of at least five different measurable signals on the screen; the amplitude of the graph (waveform) must be modifiable. The device must be equipped with an integrated printer, which in turn must meet the following requirements: the width of the printer paper 80 mm; the possibility of ECG lead printing; summary printing of the resuscitation activity; printing alarms and problems; the device must be complete with printer paper (10 pcs). The device must enable a patient’s heart to be affected by an electrical pulse defibrillation or cardioversion if necessary. Defibrillation must comply with the following requirements: application of biphasic electrical pulse; user-friendly, working on principle 1: choose energy – 2: charge – 3: shock; selection of the energy level of the electric pulse at least within the limits of 1 J to 200 J; reaching the energy level of the electric impulse when charging with a full battery 200 J in less than 7 seconds; the function of synchronising the electrical pulse with a patient’s heart rhythm. The device must allow monitoring of the heart rhythm in real-time and, if necessary, to perform extracorporeal electrocardiostimulation (electrocardiostimulator), which shall meet at least the following parameters: the stimulation current shall be changeable from 0 to 140 mA; the stimulation frequency shall be changeable from 30 to 180 beats/min. in increments of 10 beats/min.; enable the use of fixed and, if necessary, demand pacing modes. The device must be equipped with a set of batteries, meeting the following requirements: at least 6 hours of monitoring or transmission of 300 electrical impulses to the human body with an energy of at least 200 J; batteries shall be chargeable from both 230V and 12V power supply systems; automatic charging of batteries in the mains. The device must be equipped with a variable audible and visual alarm system (in compliance with IEC 60601-1-8), signalling at least the following events: battery discharge; high and low heart rate; low SpO2; high and low respiratory rate; high and low non-invasive pressure; high and low invasive pressure; life-threatening cardiac arrhythmia; low or high capnography reading. The device must meet the following general requirements: weight in full assembly up to 5.5 kg; the device must operate in accordance with all the requirements of the Estonian Defence Forces; impact and splash resistant; IEC60529 – IPX5 (water and foreign objects); MIL-STD810G – M514.6; EN ISO 9919 – IEC 60068-2.64; RTCA/DO-160G (helicopters); EN 1789 shock: IEC60068-2-27, 100 g 6 ms half sine; falling EN1789 (0.75 m), IEC 60601 (1 to 2 m). Power supply: 24 V and 230 V. | EX
3.8.9. | MUST | Gel mat set – two gel mat sets for the operating table. The set must include head, foot, shin, forearm, buttock and lumbar mats. | EX
3.8.10. | MUST | Device for heating infusion solutions – two devices with a heating temperature of 37 °C and 39 °C 38°C. Start-up time not more than 30 seconds. Capacity of at least 150 ml of liquid per minute (5 °C > 39 °C), at least 1.5 L in succession. Suitable for IV system with diameters of 4 to 5 mm or 6 to 7 mm. Heating profile length 1500 mm, Luer lock connections for attachment of IV- and blood products. Weight 1300 800g; 230V power supply connectors, 3m connection cable, IP X 2 protection rating, rechargeable, Li-ion battery. | EX
3.8.11. | MUST | Warming cabinet for infusion solutions – with a capacity of at least 20 L, power consumption of no more than 900 W, internal temperature 39 °C +/- 2 °C. | EX
3.8.12. | MUST | Refrigerator – with a capacity of at least 40 L, intended for storing medical preparations, blood and medicines. The temperature range shall be settable with an accuracy of 1 degree from +2 to +8 °C. The refrigerator must be divided into two sections with one shelf, and the shelf must withstand a load of at least 30 kg. The noise level of the appliance must be less than 50 dB(A) and the weight must be less than 60 kg. | EX
3.8.13. | MUST | Cabinet for medicines and accessories – with a capacity of at least 920 L. The cabinet shall include one 250L section with two shelves and two outswing doors; a 100L section with one outswing door; and eight 60L drawers. It must be possible to close the cabinet doors securely for transport. The exact dimensions of furniture parts to achieve the required volume will be selected during specification of the solution. | EX
3.8.14. | MUST | Secure cupboard – with a capacity of at least 7 L for the storage of medicines, integrated in the structure of the container. Lockable with a key, with the lock having at least two tumblers. Wall thickness at least 3 mm; the surface must withstand cleaning with disinfectants and antibacterial agents. | EX
3.8.15. | MUST | Service trolley with two shelves – two units, shelves and frame of stainless steel, outer dimensions of at least 650 x 500 x 800 mm, conductive wheels of at least 75 mm. | EX
3.8.16. | MUST | Mayo table – two units with adjustable height range from 95 mm to 130 mm. Tray dimensions of at least 600 x 300 mm. | EX
3.8.17. | MUST | Touch control monitor – at least 26 inches, with LED backlight and Full HD resolution, HDMI and DVI connectors, conforming to the 60601-1 medical equipment standard, and registered as a medical device. The monitor must be connected to the server computer located in the rack. | EX
3.8.18. | MUST | Anaesthetic apparatus with an anaesthetic monitor – two devices with a power of 70 W, battery-operable for 120 min., gas flow adjustment range from 0.0 to 12.0 L/min., oxygen evacuation max. 75 L/min., with electric ventilators, volume and pressure control, automatic breathing assistance, manual ventilation, spontaneous breathing support, breathing rate from 4 to 60 bpm (1/min.), PEEP 0 to 20 hPa, Pmax 15 to 70 cmH2O, tidal volume 20 to 1400 ml (20 to 1100 ml SIMV/PS), inspiratory pause (Tip:Ti) 0 to 50%, SIMV inspiration time 0.3—4.0 sec., inspiratory flow (InSPFlow) 10 to 75 L/min. (10 to 85 L/min. in PS and SIMV/PS mode), PPS PEEP +3 to 20 hPa, apnoea prevention function, trigger level 2 to 15 L/min., temperature regulation of inhalation gases, gas connections (O2, N2O, air), RS-232 data interface (possibility to transmit alarms and technical fault signals). The anaesthesia monitor system must have a response time of less than 3.0 sec.; CO2 measuring range 0 to 10 Vol %, 0 to 10 kPa, 0 to 76 mmHg; CO2 accuracy ± (0.43 Vol% + 8% rel.), ± (0.43 kPa % + 8% rel.), ± (3.3 mmHg % + 8% rel.); CO2 response time (t10/90) less than 300 ms; N2O measuring range 0 to 100 Vol%, N2O accuracy ± (2 Vol% + 8% rel.), N2O response time (t10/90) less than 300 ms. Measuring ranges for anaesthetic agents: 0 to 8.5 Vol% (isoflurane and halothane); 0 to 10 Vol% (sevoflurane and enflurane); 0 to 20 Vol% (desflurane). Accuracy on anaesthetic agents ± (0.20 Vol% + 15% rel.). Response time (t10/90) to anaesthetic agents less than 450 ms. Respiratory rate measurement range 0 to 100/min.; accuracy (I:E 1:2) 0 to 80/min.: ± 1/min., > 80/min. It must be possible to connect the anaesthetic apparatus to the anaesthetic monitor to form a single unit. The set must include a calibration-free vaporizer for sevoflurane with a working temperature range of at least 10–40 °C. | EX
3.8.19. | MUST | Medical instrument trolley with drawers – two trolleys, measuring 640 x 480 670 x 530 mm, and a worktop height of at least 950 mm. The worktop must be made of stainless steel or polyethylene, with a stainless steel frame. ISO drawer system with five drawers, measuring 600 x 400 mm. Elevated push handle. Antistatic wheels with breaks, measuring at least 125 mm. Workstation holder. Two gas cylinder holders, a 12-litre trash can, a fold-out accessory table and a small-accessory organiser (with 8–12 sections) installed above the worktop. | EX
3.8.20. | MUST | Surgical light – turbulence of not more than 27%, power not exceeding 35 W, dimming capability of central illumination 5–100%, central illumination 120 000 lx, light field diameter D10: 170–220 mm, colour rendering index (CRI): 95, battery-powered operation at least 2 h at full capacity. | EX
3.8.21. | MUST | Two magnetic wall mounts for infusion pumps/IV drip counters with a capacity of 4 devices. The wall mount must provide power and data connection to the devices. | EX
4. EMERGENCY MODULE
4.1. General
4.1.1. | MUST | It must be possible to connect the container with another container of the same type via high-pressure tunnels (corridors are for the Laboratory, connection and triage set) to form a single system. | EX
4.1.2. | MUST | The container must include an emergency exit for leaving the working space of the container when the front door is obstructed, location shall be specified during the design stage. | EX
4.1.3. | MUST | The walls and floor of the room should be easy to clean and resistant to disinfectant solutions. The resistance of the floor material must be in the range of 104–106 Ohms. | EX
4.1.4. | MUST | The container must provide automated disinfection (disinfection of all surfaces in less than 15 minutes), the device must be contained inside the container. The tenderer shall submit their own solution. | EX | A container is equipped with a UV-C air purification device NEXA PROLUX G 30WA/SPH02, and an additional autonomous and fully automatic atomizer the Anios Aerosept Ultra 15 device intended for the disinfection of previously cleaned surfaces.
4.2. Low Voltage Systems
4.2.1 | MUST | The electricity mains of the container is 3-phase 400V 50Hz. The container must be electrically autonomous, which is ensured, in addition to the possibility of connecting to the external power supply (i.e grid), by a diesel generator with a capacity of at least 8 kW. The fuel supply of the diesel generator must ensure the autonomy of the container for 36 hours at 75–80% load. The fuel tank of the container must allow refuelling from outside the container. | EX
4.2.2 | MUST | During the operation of the container's electric generator, it must be possible switch to an external power supply and back without a power interruption. In the case of power interruption, the autonomous electric generator (located inside the container) must start and connect to the network within 30 seconds. In case there is no grid connection and the operation of the generator is excluded the main systems (including lighting) of the container shall be powered by batteries, whereby the batteries shall provide at least an hour at full power. Switching with an RLA switch. | EX
4.2.3 | MUST | The standalone generator must start automatically. At startup time, uninterrupted operation of the system must be ensured (e.g. UPS). The container must have a built-in fuel tank for 36 hours of medium load operation, refueling must be done outside the container. | EX
4.2.4 | MUST | It must be possible to use the power of the autonomous electric generator to power additional containers through output connectors. | EX
4.2.5 | MUST | In addition to the 400V and 230V electrical systems, there must be 12V and 24V electrical systems if this is required for the equipment used. The container must be equipped with the required low-voltage/extra low-voltage transformers. | EX
4.2.6 | MUST | The interior power system of the container must be designed to meet the needs of the planned equipment (power outlets types and quantity, IP ratings, lighting, potential equalisation sockets, circuit breakers, protection devices etc) according to all relevant standards and be appropriate for the purpose of the container. The classification of the container in accordance with Estonian standard EVS-HD 60364-7-710 or Lithuanian standard LST HD 60364-7-710 is group 1 (as per the buyer). | EX
4.3. Extra Low Voltage (ELV) & Data Communication Systems
4.3.1 | MUST | The container must contain RJ-45 outlets (10 in total) at workstations and at least two inputs with RJF type sockets in the external feedthrough cabinet, with a possibility to close them with a cap or a stopper. The cables have an endpoint in the ELV & data communication cabinet on the RJ45 panel and they are marked. Cat6 twisted pair cable is used, the colour code of which is T568B. Cables shall be marked at both ends. | EX
4.3.2 | MUST | The ELV & data communication cabinet must be at least 10-inch, which must contain at least a 12-port Cisco (or an analogue programmable on the same principles) switch for setting up tunnels, and a Cisco (or an analogue programmable on the same principles) 4G router with 4G and a GPS antenna on the roof. The cabinet must have a server computer to run the monitoring system, generator and medical equipment software. The cabinet must be provided with proper cooling. | EX
4.4. Ilumination
4.4.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | MD
4.4.2. | MUST | Interior lighting must ensure uniform and diffused illumination, meeting the requirements for surgical premises, throughout the room. Illumination must support dimming and DALI control. The illumination of working surfaces at 950 mm from the floor must be at least 1000 lx. | EX
4.4.3. | MUST | The illumination of technical areas and rooms must be at least 200 lx if not stated otherwise. | MD
4.4.4. | MUST | The interior lightning of the container shall be ensured by the generator, external power and battaries, whereby the batteries shall provide at least an hour at full power. | EX
4.5. Ventilation
4.5.1 | MUST | The container must have proper air exchange and it must be equipped with a HEPA 13 filter. A two-stage air purifier with at least a G4 filter in the first stage and at least an F9 filter in the second stage. Mechanical intake and an electrically heated calorifier. Indoor air temperature in winter: 18–25 °C (+/-1), indoor air temperature in summer: 24 °C. | EX
4.5.2 | MUST | The air in the container must comply with ISO 7 cleanroom requirements, with UV power of 110 W, air purification rate of 140 m3/h. The room must have an overpressure of up to 20 Pa, air humidity in the range of 40–60% (reachable from 100% humidity in less than 30 minutes), air exchange at least 25 times per hour, including at least 10 times with ambient air. It must be possible to clean the volume of the container with a UV air purifier every 60 minutes. Air temperature must be selectable in the range of +18...+25 °C with a tolerance of +/-1 degrees at an outside temperature range of -35 to +40 °C, assuming that the room includes up to eight people and the specified equipment. | EX
4.6. Water and Sewerage System – None
4.7. Medical Oxygen
4.7.1. | MUST | The container must include gas pipes and a possibility of an external O2 connection to establish a network between containers. Oxygen outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. Each workstation must be provided with oxygen and compressed air outlets. For connection to the gas network, the container must include an oxygen hose of at least 10 m in length with an internal diameter of 8 mm. | EX
4.7.2. | MUST | The container must be equipped with a medical gas distribution system and a possibility for external pressurised medical air connection to create a network between containers (surgical, intensive care and emergency medical containers, triage tent). Pressurised outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. For gas network connection, the container must be provided with an external pressurised medical air inlet connection and be supplied with an oxygen hose of at least 25 m in length and an inside diameter of 8 mm. | MD
4.7.3. | MUST | On both sides of the container, there must be installed gas boxes, each containing 4 SS 8752430 or 4 DIN 13260-2 type oxygen quick-connect sockets and 2 SS 8752430 or 2 DIN 13260-2 type medical compressed air quick-connect sockets. The boxes must be interconnected and the connection must allow for the transfer of gases from one box to another without obstruction. The dimensions of the box must be chosen according to the quick-connect sockets used, but not larger than 800x600x300 mm. | MD
4.8. Equipment
4.8.1. | MUST | All medical devices must comply with Regulation (EU) 2017/745 of European parliament and of the council of 5 April 2017 on medical devices. | EX
4.8.2. | MUST | Monitor-defibrillator – two defibrillators with monitor, recording the physiological parameters of a patient, enabling one to monitor the following parameters of a patient: registration of a 12-lead ECG; registration of a 3- or 5-lead ECG; heart rate measurement with a measuring range of at least from 30 to 300 BPM; pulse oximetry (SpO2) with a measuring range of at least from 1 to 100%, to an accuracy of ± 2%; heart rate with a measuring range of at least from 25 to 240 BPM; monitoring of non-invasive blood pressure (NIBP) with a measuring range of at least from 40 to 270 mmHg; monitoring of invasive blood pressure (IBP) with a measuring range of at least from -30 to 300 mmHg; capnography monitoring using sidestream capnography equipped with Microstream accessories or equivalent technology the Microstream method; the device must provide feedback on the quality of performed chest compressions during resuscitation according to ECR/AHA guidelines; the device must enable round-the-clock registration of administered drugs and trend monitoring of measured parameters on the screen, as well as the printout of trends. The device must be complete with the following tools: a set of cables and electrodes for recording a 12-lead ECG; cable and sensor for monitoring pulse oximetry; different sized cuffs and hoses for non-invasive blood pressure monitoring; invasive blood pressure monitoring kit; capnography sensor; a wall mount; charger; with a protective case and closable accessory pockets (cables, electrodes, sensors, etc.) and a shoulder strap. The device screen must meet the following requirements: a diagonal screen of at least 16 cm, coloured; allow simultaneous monitoring of at least five different measurable signals on the screen; the amplitude of the graph (waveform) must be modifiable. The device must be equipped with an integrated printer, which in turn must meet the following requirements: the width of the printer paper 80 mm; the possibility of ECG lead printing; summary printing of the resuscitation activity; printing alarms and problems; the device must be complete with printer paper (10 pcs). The device must enable a patient’s heart to be affected by an electrical pulse defibrillation or cardioversion if necessary. Defibrillation must comply with the following requirements: application of biphasic electrical pulse; user-friendly, working on principle 1: choose energy – 2: charge – 3: shock; selection of the energy level of the electric pulse at least within the limits of 1 J to 200 J; reaching the energy level of the electric impulse when charging with a full battery 200 J in less than 7 seconds; the function of synchronising the electrical pulse with a patient’s heart rhythm. The device must allow monitoring of the heart rhythm in real-time and, if necessary, to perform extracorporeal electrocardiostimulation (electrocardiostimulator), which shall meet at least the following parameters: the stimulation current shall be changeable from 0 to 140 mA; the stimulation frequency shall be changeable from 30 to 180 beats/min. in increments of 10 beats/min.; enable the use of fixed and, if necessary, demand pacing modes. The device must be equipped with a set of batteries, meeting the following requirements: at least 6 hours of monitoring or transmission of 300 electrical impulses to the human body with an energy of at least 200 J; batteries shall be chargeable from both 230V and 12V power supply systems; automatic charging of batteries in the mains. The device must be equipped with a variable audible and visual alarm system (in compliance with IEC 60601-1-8), signalling at least the following events: battery discharge; high and low heart rate; low SpO2; high and low respiratory rate; high and low non-invasive pressure; high and low invasive pressure; life-threatening cardiac arrhythmia; low or high capnography reading. The device must meet the following general requirements: weight in full assembly up to 5.5 kg; the device must operate in accordance with all the requirements of the Estonian Defence Forces; impact and splash resistant; IEC60529 – IPX5 (water and foreign objects); MIL-STD810G – M514.6; EN ISO 9919 – IEC 60068-2.64; RTCA/DO-160G (helicopters); EN 1789 shock: IEC60068-2-27, 100 g 6 ms half sine; falling EN1789 (0.75 m), IEC 60601 (1 to 2 m). Power supply: 24 V and 230 V. | EX
4.8.3. | MUST | Aspirators – two devices: air flow rate at least 45 L/min., maximum vacuum at sea level -91 90 kPa- 900 mbar- 675 mmHg under normal conditions, vacuum level accuracy: -1–0 bar ± 16 mbar (Class 1.6), possibility to use 2 x 3L or 2 x 5L 2L/4L/5L transparent secretion tanks, possibility to use ø 6mm 8mm or ø 10mm hoses, maximum power not more than 200 W, weight with secretion tank not more than 24 kg, noise level < 48 dB (A) @ 1 m (ISO 7779). | EX
4.8.4. | MUST | Oxygen generator – the oxygen generator must produce oxygen at the rate of at least 20 L/oxygen per minute with a pressure of 5 bars and an oxygen concentration of 93%. The weight of the oxygen generator must not exceed 120 kg. The current consumption shall not exceed 1.4 kVA. The oxygen generator must be single-phase. The oxygen generator must comply with the IP rating of 42. The oxygen generator must have three oxygen outlets, including at least one with adjustable pressure between 0 and 8 bars; one outlet must support flow rate adjustment and one outlet must be designed to connect the device to an oxygen bottling system. The current oxygen pressure must be indicated by a manometer. The generator must have a separate outlet for discharging oxygen-depleted air. It must be possible to remove the oxygen generator from the container if necessary. The oxygen generator must be supplied with a packing box or boxes with MIL-15 certification and IP65 protection rating. The oxygen generator must be provided with a connecting hose for connecting the generator to the DIN 13260-2 or SS 8752430 standard quick coupling. A 10 L replaceable oxygen cylinder connected to a compressor in a way that the oxygen supply continues automatically from the cylinder in case the compressor stops working. | EX
4.8.5. | MUST | Medical air compressor – with an output capacity of at least 40 LPM, output pressure up to 7 bar, noise level not exceeding 53 dBA. Air purity at least Class 0 (ISO 8573-1:2010). Continuous air production at least 100 LPM, configurable from 0 to 10 bar. Connected to pipework stipulated in clause 4.7.2. | EX
4.8.6. | MUST | Respiratory ventilator – two portable respiratory ventilators suitable for use on patients weighing 5 kg or more. The devices shall allow a patient to be ventilated at least in AC, SIMV, CPAP, and BL modes. The device must have an LCD screen that displays at least: waveform, battery status, selected mode, HR, SpO2, FiO2, PIP, Vt, BPM. The device must allow the use of both pressure- and volume-controlled modes and ventilation with at least the following parameters: frequency at least from 1 to 80 60 bpm, breathing volume at least from 50 to 2000 ml, PEEP from 0 to 30 mbar, 20cm H2O inspiratory time from 0.3 to 5 4 sec, I:E ratio at least from 1:99 4:1 to 4:1, O2 concentration from 21 to 100%, trigger sensitivity variable from -0.5 to 6.0 cmH2O. The device must be equipped with at least the following alarms: high airway pressure, adjustable between 20 and 100 cmH2O; low airway pressure, which must be with the option of being switched off; high breathing rate, variable between 2 and 99 12 and 70 BPM; low breathing rate. The device must have at least a three-step visual alarm indicator. The weight of the device must not exceed 4.5 kg. The operating time of the integrated battery must be at least 10 hours, a battery charge of up to 90% must not require more than 2 hours. Operating time 4 hours, charging time 4 hours. The device must be able to be used in air transport at least at an altitude of up to 7620 m, must be AWR (Airworthiness Release) certified. The dimensions of the device must not exceed 205 x 320 x 115 mm. Must comply with MIL STD 810F standard. It must be possible to mute the device and put it in dark mode. | EX
4.8.7. | MUST | Infusion pump – eight infusion pumps with dimensions not exceeding 245 x 90 x 180 mm, easy to use, mobile and easy to clean, weighing not more than 2.0 kg with an integrated fixation clip for the drip stand and a battery. Mains power 230 V, and the device must operate from both the mains and the lithium battery, with a battery operating time of at least 10 h 6 h at an infusion rate of 5 ml/h. The device must allow for the use of different syringes (different types of sizes 5, 10, 20, 30, 50/60 ml, recognised automatically by the infusion pump) and functions: set volume of infusion; specified infusion duration; administration of a bolus; the infusion pump must have an anti-bolus function. The device must have a coloured touch screen, indicating at least the following parameters: infusion rate (ml/h); the amount of infused liquid (at least from 0.1 to 999.9 ml); the time until the infusion is complete/the syringe emptied; occlusion level (settable within the range of approximately 300 to 900 mbar or equivalent from 0 to 1000 mbar and easily modifiable); event trend, at least 1500 per an event. The parameters must be changeable at least in the following range: rate from 0.01 to 1500 ml/h (for 10–50ml syringes); set volume of infusion from 0.1 to 100 ml; infusion duration from 1 min. to 24 hours; a keyboard shortcut must be provided to fill a bolus. The accuracy of the infusion must be at least ± 2% in accordance with the IEC/EN 60601-2-24 standards. The alarms must have both two lights representing two levels and a sound announcing at least the end of the infusion, reaching the set volume of infusion, remaining time until the end of the infusion (a few minutes prior), battery discharge, battery capacity, 230V external mains failure, technical error. The device must be equipped with audible and visual alarms indicating at least infusion completion, reaching the set infusion volume, low battery level, external mains failure, occlusion, and technical errors. Possibility to switch the syringe pumps to the alarm system (nurse call). The instructional text transmitted on the pump screen must be in Estonian or Lithuanian or English. Infrared port, in addition to data transfer, also for inserting new software upgrades. The device must be equipped with an infrared port, USB port, or equivalent interface for data transfer and software upgrades. The syringe plunger must not protrude beyond the syringe pump housing in the case of a filled syringe. Possibility to connect syringe pumps to a special quick-coupled storage compartment with power supply. | EX
4.8.8. | MUST | Device for heating infusion solutions – two devices with a heating temperature of 37 °C and 39 °C 38°C. Start-up time not more than 30 seconds. Capacity of at least 150 ml of liquid per minute (5 °C > 39 °C), at least 1.5 L in succession. Suitable for IV system with diameters of 4 to 5 mm or 6 to 7 mm. Heating profile length 1500 mm, Luer lock connections for attachment of IV- and blood products. Weight 1300 800g; 230V power supply connectors, 3m connection cable, IP X 2 protection rating, rechargeable, Li-ion battery. | EX
4.8.9. | MUST | Warming cabinet for infusion solutions – with a capacity of at least 20 L. Power consumption of no more than 900 W, internal temperature 39 °C +/- 2 °C. | EX
4.8.10. | MUST | Refrigerator – for medicines and solutions, with a capacity of at least 40 L, intended for storing medical preparations and medicines. It must be possible to set the temperature with an accuracy of 1 degree in the range from +2 to +8 °C. The refrigerator must be divided into two sections with one shelf, and the shelf must withstand a load of at least 30 kg. The noise level of the appliance must be less than 50 dB(A) and the weight must be less than 60 kg. | EX
4.8.11. | MUST | Cabinet for medicines and accessories – with a capacity of at least 920 L. The cabinet shall include one 250L section with two shelves and two outswing doors; a 100L section with one outswing door; and eight 60L drawers. It must be possible to close the cabinet doors securely for transport. The exact dimensions of furniture parts to achieve the required volume will be selected during specification of the solution. | EX
4.8.12. | MUST | Secure cupboard – with a capacity of at least 7 L, integrated in the structure of the container. Lockable with a key; the surface must withstand cleaning with disinfectants and antibacterial agents. | EX
4.8.13. | MUST | Service trolley – two service trolleys with two shelves. Shelves and frame made of stainless steel, measuring at least 650 x 500 x 800 mm, conductive wheels of at least 75 mm. | EX
4.8.14. | MUST | Trolley for medical instruments – two trolleys with drawers for medical instruments. Dimensions: 640 x 480 670 x 530 mm; worktop height at least 1000 950 mm. Worktop made of material stainless steel or polyethylene. Stainless steel frame. ISO drawer system. Drawers measuring 600 x 400 mm. Five drawers in total. Elevated push handle. Antistatic wheels with breaks, measuring at least 125 mm. Workstation holder. Two gas cylinder holders, a 12L trash can, a fold-out accessory table and a small-accessory organiser (with 8–12 sections) installed above the worktop. | EX
4.8.15. | MUST | Mobile X-ray machine with a digital receptor plate – X-ray must be configurable in 1 kVp increments between 40 and 125 kVp. Configurable mA range from 5 to 100 mA. Exposure time can be set between 0.001–10 seconds. Double-focus X-ray tube with a collimator and a TAP meter. Weight should not exceed 80 kg. The machine must be easy to move and easy to handle; it must be possible to move it around inside the container as well as in tents erected on soft terrain. The X-ray machine must support DICOM, HIS and RIS interfaces. The energy range of the wireless digital receptor board must be 40–150 kVp. Resolution 3.7 lp/mm. A diagnostic touch screen of at least 17 inches for viewing images. | EX
4.8.16. | MUST | Operating table – two operating tables, must be integrated in the container. The operating table must be at least 2100 x 560 mm in length, with a load capacity of at least 320 kg. The table height must be adjustable within a range of at least 680–1120 mm, with lateral tilt in both directions from -35° to +75°, and Trendelenburg and reverse Trendelenburg positions of at least 40°. The backrest must be adjustable from -35° to +75°, the leg section from -90° to +20° with a 180° rotation. The table must have a resuscitation function, i.e. the table top can be rapidly lowered to the zero position. The table top must be divisible into at least four parts (removable headrest, backrest, and leg sections). The table top must be compatible with X-ray imaging. All elements of the operating table must have accessory rails. The operating table must be equipped with trauma surgery (damage control surgery) accessories. The frame of the operating table must be made of stainless steel. The table must be disconnected from the electrical network during the procedure. Lifting, lowering, and positioning must be done electrically and the table must have a built-in battery. The table must be grounded or double insulated. The table must be able to be connected to the electrical network during the operation or operated on battery power. A potential equalization cable of at least 5 m must be supplied with the table. An operating lamp must be provided above the table. | EX
4.8.17. | MUST | Patient warming device – two electric patient warming blowers with a heating power of at least 460 W and an average power consumption of no more than 800 W. The temperature of the blown air must be adjustable to values of 32 °C, 38 °C, and 43 °C. The device must reach the required temperature within a maximum of 5 minutes. The device must have an overheat alarm. The device must come with consumables for 100 patients. The consumables must include a lower body, upper body, and patient base warming blanket. The device must weigh no more than 8 kg and produce no more than 54 55dB of noise. | EX
4.8.18. | MUST | Oxygen humidifier – two oxygen humidifiers with heating function. Weight not more than 2.5kg. Target temperature values: 31 °C, 34 °C, 37°C. Heating rate of 10 minutes at 31 °C and 30 minutes at 37 °C. Moisture output of at least 32mg/L at 37 °C and at least 10mg/L at 31 °C. Maximum oxygen flow rate of 60 L/min. The device must have an oxygen sensor with an accuracy of ± 4%. The device must come with a disinfection kit. The device must come with a cart or appropriate magnetic wall mounts for the entire device and its accessories. The device must be maintenance-free. | EX
4.8.19. | MUST | Two magnetic wall mounts for infusion pumps/IV drips with a capacity of 4 devices. The wall mount must provide power and data connection to the devices. | EX
5. INTENSIVE CARE MODULE
5.1. General
5.1.1. | MUST | It must be possible to connect the container with another container of the same type via high-pressure corridor to form a single system. | EX
5.1.2. | MUST | The container has to be supplied with a high-pressure corridor that can be connected with the door ends of the container (on two sides). The high-pressure corridors shall provide the possibility of connecting the container to other containers. | EX
5.1.3. | MUST | Max. pressure in the high pressure corridor may not exceed 8 bar. The high-pressure corridor must include removable and washable fabric used for insulation. | EX
5.1.4. | MUST | There must be a main entrance door at one end of the container and an exit door at the other end, high-pressure corridors must be attached to the doors/exits. | EX
5.1.5. | MUST | It must be possible to walk through the container, i.e. it may not include partitions or any fixed equipment installed in the passageway. | EX
5.1.6. | MUST | The walls and floor of the room should be easy to clean and resistant to disinfectant solutions. The resistance of the floor material must be in the range of 104–106 Ohms. | EX
5.1.7. | MUST | The container must provide automated disinfection (disinfection of all surfaces in less than 15 minutes), the device must be contained inside the container. The tenderer shall submit their own solution. | EX | A container is equipped with a UV-C air purification device NEXA PROLUX G 30WA/SPH02, and an additional autonomous and fully automatic atomizer the Anios Aerosept Ultra 15 device intended for the disinfection of previously cleaned surfaces.
5.2. Low Voltage Systems
5.2.1 | MUST | The electricity mains of the container is 3-phase 400V 50Hz. The container must be powered from an external supply (grid or generator). In case there is no external supply for the generator the main systems (including lighting) of the container shall be powered by batteries, whereby the batteries shall provide at least an hour at full power. | EX
5.2.2 | MUST | In addition to the 400V and 230V electrical systems, there must be 12V and 24V electrical systems if this is required for the equipment used. The container must be equipped with the required low-voltage/extra low-voltage transformers. | EX
5.2.3 | MUST | The interior power system of the container must be designed to meet the needs of the planned equipment (power outlets types and quantity, IP ratings, lighting, potential equalisation sockets, circuit breakers, protection devices etc) according to all relevant standards and be appropriate for the purpose of the container. The classification of the container in accordance with Estonian standard EVS-HD 60364-7-710 and Lithuanian standard LST HD 60364-7-710 is group1 (as per the buyer). | EX
5.3. Extra Low Voltage (ELV) & Data Communication Systems
5.3.1. | MUST | The container must contain RJ-45 outlets (14 in total) at workstations and at least two inputs with RJF type sockets in the external feedthrough cabinet, with a possibility to close them with a cap or a stopper. The cables have an endpoint in the ELV & data communication cabinet on the RJ45 sockets on the respective panel, and must be marked. Cat6 twisted pair cable is used, the colour code of which is T568B. Cables are marked at both ends. | EX
5.3.2. | MUST | The ELV & data communication cabinet must be at least 10-inch, which must contain at least a 16-port Cisco switch (or equivalent, programmable according to the same principles) for setting up tunnels, and a Cisco 4G router (or equivalent, programmable according to the same principles) with 4G and a GPS antenna on the roof. The cabinet must have a server computer to run the monitoring system, generator and medical equipment software. The cabinet must be provided with proper cooling. | EX
5.4. Ilumination
5.4.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | MD
5.4.2. | MUST | Interior lighting must ensure uniform and diffused illumination, meeting the requirements for surgical premises, throughout the room. Illumination must support dimming and DALI control. The illumination of working surfaces at 950 mm from the floor must be at least 1000 lx. | EX
5.4.3. | MUST | The illumination of technical areas and rooms must be at least 200 lx if not stated otherwise. | MD
5.4.4. | MUST | The interior lightning of the container shall be ensured by the generator, external power and battaries, whereby the batteries shall provide at least an hour at full power. | EX
5.5. Ventilation
5.5.1. | MUST | The container must have proper air exchange and it must be equipped with a HEPA 13 filter. Two-stage air purifier, with the first stage at least G4 and the second at least F9. Mechanical intake and electric heating with a calorifier. Indoor air temperature in winter: 18–25 °C (±2), indoor air temperature in summer: 24 °C. | EX
5.5.2. | MUST | The air in the container must comply with ISO 7 cleanroom requirements, with UV power of 55 W, air purification rate of 140 m3/h. The room must have an overpressure of up to 20 Pa, air humidity in the range of 30–60% (reachable from 100% humidity in less than 30 minutes), air exchange at least 15 times per hour, including at least eight times with ambient air. It must be possible to clean the volume of the container with a UV air purifier every 60 minutes. Air temperature must be selectable in the range of +18...+25 °C with a tolerance of +/-1 degrees at an outside temperature range of -35 to +40 °C, assuming that the room includes up to eight people and the specified equipment. | EX
5.6. Water Supply and Sewerage System – None
5.7. Medical Oxygen
5.7.1 | MUST | The container must include gas pipes and a possibility of an external O2 connection for the connection of gases between containers. Oxygen outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. Each workstation must be provided with oxygen and compressed air outlets. For connection to the gas network, the container must include an oxygen hose of at least 10 m in length with an internal diameter of 8 mm. | EX
5.7.2 | MUST | The container must be equipped with a medical gas distribution system and a possibility for external pressurised medical air connection to create a network between containers (surgical, intensive care and emergency medical containers, triage tent). Pressurised outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. For gas network connection, the container must be provided with an external pressurised medical air inlet connection and be supplied with an oxygen hose of at least 25 m in length and an inside diameter of 8 mm. | MD
5.7.3 | MUST | On both sides of the container, there must be installed gas boxes, each containing 4 SS 8752430 or 4 DIN 13260-2 type oxygen quick-connect sockets and 2 SS 8752430 or 2 DIN 13260-2 type medical compressed air quick-connect sockets. The boxes must be interconnected and the connection must allow for the transfer of gases from one box to another without obstruction. The dimensions of the box must be chosen according to the quick-connect sockets used, but not larger than 800x600x300 mm. | MD
5.8.Equipment
5.8.1. | MUST | All medical devices must comply with Regulation (EU) 2017/745 of European parliament and of the council of 5 April 2017 on medical devices. | EX
5.8.2. | MUST | Monitor-defibrillator – four defibrillators with monitor, recording the physiological parameters of a patient, enabling one to monitor the following parameters of a patient: registration of a 12-lead ECG; registration of a 3- or 5-lead ECG; heart rate measurement with a measuring range of at least from 30 to 300 BPM; pulse oximetry (SpO2) with a measuring range of at least from 1 to 100%, to an accuracy of ± 2%; heart rate with a measuring range of at least from 25 to 240 BPM; monitoring of non-invasive blood pressure (NIBP) with a measuring range of at least from 40 to 270 mmHg; monitoring of invasive blood pressure (IBP) with a measuring range of at least from -30 to 300 mmHg; capnography monitoring using sidestream capnography equipped with Microstream accessories or equivalent technology the Microstream method; the device must provide feedback on the quality of performed chest compressions during resuscitation according to ECR/AHA guidelines; the device must enable round-the-clock registration of administered drugs and trend monitoring of measured parameters on the screen, as well as the printout of trends. The device must be complete with the following tools: a set of cables and electrodes for recording a 12-lead ECG; cable and sensor for monitoring pulse oximetry; different sized cuffs and hoses for non-invasive blood pressure monitoring; invasive blood pressure monitoring kit; capnography sensor; a wall mount; charger; with a protective case and closable accessory pockets (cables, electrodes, sensors, etc.) and a shoulder strap. The device screen must meet the following requirements: a diagonal screen of at least 16 cm, coloured; allow simultaneous monitoring of at least five different measurable signals on the screen; the amplitude of the graph (waveform) must be modifiable. The device must be equipped with an integrated printer, which in turn must meet the following requirements: the width of the printer paper 80 mm; the possibility of ECG lead printing; summary printing of the resuscitation activity; printing alarms and problems; the device must be complete with printer paper (10 pcs). The device must enable a patient’s heart to be affected by an electrical pulse defibrillation or cardioversion if necessary. Defibrillation must comply with the following requirements: application of biphasic electrical pulse; user-friendly, working on principle 1: choose energy – 2: charge – 3: shock; selection of the energy level of the electric pulse at least within the limits of 1 J to 200 J; reaching the energy level of the electric impulse when charging with a full battery 200 J in less than 7 seconds; the function of synchronising the electrical pulse with a patient’s heart rhythm. The device must allow monitoring of the heart rhythm in real-time and, if necessary, to perform extracorporeal electrocardiostimulation (electrocardiostimulator), which shall meet at least the following parameters: the stimulation current shall be changeable from 0 to 140 mA; the stimulation frequency shall be changeable from 30 to 180 beats/min. in increments of 10 beats/min.; enable the use of fixed and, if necessary, demand pacing modes. The device must be equipped with a set of batteries, meeting the following requirements: at least 6 hours of monitoring or transmission of 300 electrical impulses to the human body with an energy of at least 200 J; batteries shall be chargeable from both 230V and 12V power supply systems; automatic charging of batteries in the mains. The device must be equipped with a variable audible and visual alarm system (in compliance with IEC 60601-1-8), signalling at least the following events: battery discharge; high and low heart rate; low SpO2; high and low respiratory rate; high and low non-invasive pressure; high and low invasive pressure; life-threatening cardiac arrhythmia; low or high capnography reading. The device must meet the following general requirements: weight in full assembly up to 5.5 kg; the device must operate in accordance with all the requirements of the Estonian Defence Forces; impact and splash resistant; IEC60529 – IPX5 (water and foreign objects); MIL-STD810G – M514.6; EN ISO 9919 – IEC 60068-2.64; RTCA/DO-160G (helicopters); EN 1789 shock: IEC60068-2-27, 100 g 6 ms half sine; falling EN1789 (0.75 m), IEC 60601 (1 to 2 m). Power supply: 24 V and 230 V. | EX
5.8.3. | MUST | Respiratory ventilator – four portable respiratory ventilators suitable for use on patients weighing 5 kg or more. The devices shall allow a patient to be ventilated at least in AC, SIMV, CPAP, and BL modes. The device must have an LCD screen that displays at least: waveform, battery status, selected mode, HR, SpO2, FiO2, PIP, Vt, BPM. The device must allow the use of both pressure- and volume-controlled modes and ventilation with at least the following parameters: frequency at least from 1 to 80 bpm, breathing volume at least from 50 to 2000 ml, PEEP from 0 to 30 mbar, inspiratory time from 0.3 to 5 sec, I:E ratio at least from 1:99 to 4:1, O2 concentration from 21 to 100%, trigger sensitivity variable from -0.5 to 6.0 cmH2O. The device must be equipped with at least the following alarms: high airway pressure, adjustable between 20 and 100 cmH2O; low airway pressure, which must be with the option of being switched off; high breathing rate, variable between 2 and 99 BPM; low breathing rate. The device must have at least a three-step visual alarm indicator. The weight of the device must not exceed 4.5 kg. The operating time of the integrated battery must be at least 10 hours, a battery charge of up to 90% must not require more than 2 hours. The device must be able to be used in air transport at least at an altitude of up to 7620 m, must be AWR (Airworthiness Release) certified. The dimensions of the device must not exceed 205 x 320 x 115 mm. Must comply with MIL STD 810F standard. It must be possible to mute the device and put it in dark mode. | EX
5.8.4. | MUST | Device for heating infusion solutions – four devices for heating infusion solutions, with heating temperatures of 37 °C and 39 °C; with a start-up time of 30 seconds; a capacity of at least 150 ml of liquid per minute (5 °C > 39 °C), at least 1.5 L in succession. Suitable for IV system with diameters of 4 to 5 mm or 6 to 7 mm. The length of the heating profile shall be 1500 mm; device weight 1300 g; 230V and 12V power supply connections, 3m connection cable, IP X 2 protection rating. Warming cabinet for infusion solutions with a capacity of at least 20 L. Power consumption of no more than 900 W, internal temperature 39 °C +/- 2 °C. | EX
5.8.5. | MUST | Refrigerator – for medicines and solutions, with a capacity of at least 40 L, intended for storing medical preparations and medicines. It must be possible to set the temperature with an accuracy of 1 degree in the range from +2 to +8 °C. The refrigerator must be divided into two sections with one shelf, and the shelf must withstand a load of at least 30 kg. The noise level of the appliance must be less than 50 dB(A) and the weight must be less than 60 kg. | EX
5.8.6. | MUST | Intensive care bed – four folding intensive care beds for a field hospital. NSN6530-01-624-7626. | EX
5.8.7. | MUST | Procedure lights – four procedure lights with 35,000 lx power and CRI>=90, mounted on the wall with magnets. The lifespan of the light source is at least 50 000 hours. The arm length is at least 1 meter, and the weight of the light fixture is not more than 5 kg. | EX
5.8.8. | MUST | Cabinet on wheels – for the refrigerator and the warming cabinet. | EX
5.8.9. | MUST | Equipment trolley – two equipment trolleys for mounting monitors. The trolley must also accommodate at least two suspended bags of infusion solution and the attachment of at least two infusion pumps and/or drip counters. The equipment trolley must include a shelf or drawer for accessories. The trolley must be stable and it must be possible to attach the trolley to the container for transportation. | EX
5.8.10. | MUST | Infusion pump – ten infusion pumps with dimensions not exceeding 245 x 90 x 180 mm, easy to use, mobile and easy to clean, weighing not more than 2.0 kg with an integrated fixation clip for the drip stand and a battery. Mains power 230 V, and the device must operate from both the mains and the lithium battery, with a battery operating time of at least 10 h 6 h at an infusion rate of 5 ml/h. The device must allow for the use of different syringes (different types of sizes 5, 10, 20, 30, 50/60 ml, recognised automatically by the infusion pump) and functions: set volume of infusion; specified infusion duration; administration of a bolus; the infusion pump must have an anti-bolus function. The device must have a coloured touch screen, indicating at least the following parameters: infusion rate (ml/h); the amount of infused liquid (at least from 0.1 to 999.9 ml); the time until the infusion is complete/the syringe emptied; occlusion level (settable within the range of approximately 300 to 900 mbar or equivalent from 0 to 1000 mbar and easily modifiable); event trend, at least 1500 per an event. The parameters must be changeable at least in the following range: rate from 0.01 to 1500 ml/h (for 10–50ml syringes); set volume of infusion from 0.1 to 100 ml; infusion duration from 1 min. to 24 hours; a keyboard shortcut must be provided to fill a bolus. The accuracy of the infusion must be at least ± 2% in accordance with the IEC/EN 60601-2-24 standards. The alarms must have both two lights representing two levels and a sound announcing at least the end of the infusion, reaching the set volume of infusion, remaining time until the end of the infusion (a few minutes prior), battery discharge, battery capacity, 230V external mains failure, technical error. The device must be equipped with audible and visual alarms indicating at least infusion completion, reaching the set infusion volume, low battery level, external mains failure, occlusion, and technical errors. Possibility to switch the syringe pumps to the alarm system (nurse call). The instructional text transmitted on the pump screen must be in Estonian or Lithuanian or English. Infrared port, in addition to data transfer, also for inserting new software upgrades. The device must be equipped with an infrared port, USB port, or equivalent interface for data transfer and software upgrades. The syringe plunger must not protrude beyond the syringe pump housing in the case of a filled syringe. Possibility to connect syringe pumps to a special quick-coupled storage compartment with power supply. | EX
5.8.11. | MUST | Aspirator – four aspirators with an aspiration capacity of 34 L/min. (± 4 L/min.); battery life of at least 60 min.; charging time of up to 4 h; empty weight of up to 4 kg; weight including the secretion container of up to 5 kg; noise level below 60 dB (A). The size of the secretion container must be at least 1000 ml. Power supply 230V AC and 12V or 24V DC. The device must be wall-mountable. | EX
5.8.12. | MUST | Infusor/drip counter - four infusors/drip counters. With an Estonian or Lithuanian or English language user interface, flow rate of 1...500 ml/h, infusion volume of 0.1...999.9 ml. Lithium-ion battery with a minimum runtime of 12 hours, occlusion limit of 0...999 mmHg with bolus reduction function. | EX
5.8.13. | MUST | Trolley for medical instruments – two trolleys with drawers for medical instruments. Dimensions: 640 x 480 670 x 530 mm; worktop height at least 1000 950 mm. Worktop made of material stainless steel or polyethylene. Stainless steel frame. ISO drawer system. Drawers measuring 600 x 400 mm. Five drawers in total. Elevated push handle. Antistatic wheels with breaks, measuring at least 125 mm. Workstation holder. Two gas cylinder holders, a 12L trash can, a fold-out accessory table and a small-accessory organiser (with 8–12 sections) installed above the worktop. | EX
5.8.14. | MUST | Patient warming device – two electric patient warming blowers with a heating power of at least 460 W and an average power consumption of no more than 800 W. The temperature of the blown air must be adjustable to values of 32 °C, 38 °C, and 43 °C. The device must reach the required temperature within a maximum of 5 minutes. The device must have an overheat alarm. The device must come with consumables for 100 patients. The consumables must include a lower body, upper body, and patient base warming blanket. The device must weigh no more than 8 kg and produce no more than 54 55dB of noise. | EX
5.8.15. | MUST | Oxygen humidifier – two oxygen humidifiers with heating function. The weight should not exceed 2.5 kg. Temperature target values: 31 °C, 34 °C, 37 °C. The heating rate is 10 minutes at 31 °C and 30 minutes at 37 °C. Moisture output is at least 32 mg/l at 37 °C and at least 10 mg/l at 31 °C. The maximum oxygen flow rate is 60 l/min. The device must have an oxygen sensor with an accuracy of ± 4%. The device must come with a disinfection kit. The device must have a device cart or corresponding magnetic wall mounts for the entire device and its accessories. The device must be maintenance-free. | EX
5.8.16. | MUST | Four magnetic wall mounts for infusion pumps/IV drip sets with a capacity of 4 devices. The wall mounts must provide power and data connectivity to the devices. | EX
6. TREATMENT MODULE
6.1. General
6.1.1. | MUST | The system shall comprise a two-in-one container and two high-pressure tents (width min. 6 m, floor area at least 48 m2). The tent is planned to be used for 10 patient spaces, and three high-pressure corridors connecting the continer and tents. The container and tents must have a shared power supply and communication system. | EX
6.1.2. | MUST | The container must be capable of being connected to another container of the same type and to the tents of the set, to allow for the creation of a single system. | EX
6.1.3. | MUST | Both ends and the non-opening side of the container should be provided with lowerable platforms, measuring 1500 x 2100 mm. In the transport position, the platform door is closed and it must be weatherproof. | EX
6.1.4. | MUST | The doorway must be designed and constructed in such a way that a connecting corridor with a tent can be attached to the outside. | EX
6.1.5. | MUST | It must be possible to walk through the container, i.e. it may not include partitions or any fixed equipment installed in the passageway. | EX
6.1.6. | MUST | The container must include integrated furniture made of stainless steel, including two folding tables for workstations. Two-door and single-door stainless steel cabinets with a volume of up to 700 L; the exact solution will be coordinated with the contracting authority during the design process. | EX
6.1.7. | MUST | The high-pressure tent must have two doors, one at each end, and a side door; the dimensions and the position of the doors will be coordinated with the contracting authority durin design process. | EX
6.1.8. | MUST | The high-pressure tent must include at least three partitions for dividing the tent into different zones. The tenderer shall submit their own solution. | EX | Tent is equipped with 3 separating walls that can be attached to the tent inner liner with velcro or carabiners
6.1.9. | MUST | The set must include all the equipment required for handling the container and the tents. The exact solution shall be specified during the design stage. | EX
6.2. Low Voltage Systems
6.2.1. | MUST | The electricity mains of the container is 3-phase 400V 50Hz. The container must be powered from an external supply (grid or generator). In case there is no external supply for the generator the main systems (including lighting) of the container shall be powered by batteries, whereby the batteries shall provide at least an hour at full power. | EX
6.2.2. | MUST | In addition to the 400V and 230V electrical systems, there must be 12V and 24V electrical systems if this is required for the equipment used. The container must be equipped with the required low-voltage/extra low-voltage transformers. | EX
6.2.3. | MUST | The electrical system of the container must be designed and constructed to provide 3-phase 32A 400V and 1-phase 16A 230V outlets to the two high-pressure tents included in the kit. | EX
6.2.4. | MUST | The interior power system of the container must be designed to meet the needs of the planned equipment (power outlets types and quantity, IP ratings, lighting, potential equalisation sockets, circuit breakers, protection devices etc) according to all relevant standards and be appropriate for the purpose of the container. | EX
6.3. Extra Low Voltage (ELV) & Data Communication Systems
6.3.1. | MUST | The container must contain RJ45 outlets (eight in total) at workstations and at least two inputs with RJF type sockets in the external feedthrough cabinet, with a possibility to close them with a cap or a stopper. The cables have an endpoint in the ELV & data communication cabinet on the RJ45 panel and they are marked. Cat6 twisted pair cable is used, the colour code of which is T568B. Cables are marked at both ends. | EX
6.3.2. | MUST | The ELV & data communication cabinet must be at least 10-inch, which must contain at least a 12-port Cisco (or an analogue programmable on the same principles) switch for setting up tunnels, and a Cisco (or an analogue programmable on the same principles) 4G router with 4G and a GPS antenna on the roof. The cabinet must have a server computer to run the monitoring system, generator and medical equipment software. The cabinet must be provided with proper cooling. | EX
6.4. Illumination
6.4.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | EX
6.4.2. | MUST | Interior lighting shall provide uniform and diffused illumination throughout the room. Illumination must support dimming and DALI control. The illumination of working surfaces at 950 mm from the floor must be at least 1000 lx. | EX
6.4.3. | MUST | The container must be fitted with a telescopic mast and LED spotlights for exterior lighting. The exact solution will be coordinated with the contracting authority during the design phase. | EX
6.4.4. | MUST | The illumination of technical areas and rooms must be at least 200 lx if not stated otherwise. | EX
6.4.5. | MUST | The interior lightning of the container shall be ensured by the generator, external power and battaries, whereby the batteries shall provide at least an hour at full power. | EX
6.5. Ventilation
6.5.1. | MUST | The container must exchange air at least 4 times per hour. Air temperature must be selectable in the range of +18...+25 °C with a tolerance of +/-2 degrees at an outside temperature range of -35 to +40 °C, assuming that the room includes up to eight people and the specified equipment. Air intake with heat exchanger of at least 30 l/s. | EX
6.6. Water Supply and Sewerage System
6.6.1. | MUST | The container shall contain a two-tap hand-washing station measuring at least 580 mm in width, not more then 200 mm in length and at least 100 mm in depth. A handwashing station with a stainless-steel sink and rounded internal corners. The container for clean water should have a capacity of at least 50 L. The system must have a boiler with a capacity of at least 10 L and a power of 2 kW. The handwashing station must have a pump to draw clean water into the container and create water pressure in the system. To pump clean water into the container, a water pipe must be led to the external surface of the container, through which the container can be connected to an external water source permanently or temporarily. The used water flows out through a pipe led to the external surface of another container, which can be connected temporarily or permanently to a sewage system or used water tank. The connection to the clean water source and sewage system must be possible independently of the season. | EX
6.6.2. | MUST | The handwashing station must be integrated with the furniture installed in the container. | EX
6.7. Medical Gases
6.7.1. | MUST | The container must include gas pipes and a possibility of an external O2 connection to establish a network between containers. Oxygen outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. For connection to the gas network, the container must include an oxygen hose of at least 25 m in length with an internal diameter of 8 mm. | EX
6.7.2. | MUST | The container must be equipped with a medical gas distribution system and a possibility for external pressurised medical air connection to create a network between containers (surgical, intensive care and emergency medical containers, triage tent). Pressurised outlet connectors must comply with the DIN 13260-2 or SS 8752430 standard. For gas network connection, the container must be provided with an external pressurised medical air inlet connection and be supplied with an oxygen hose of at least 25 m in length and an inside diameter of 8 mm. | EX
6.7.3. | MUST | On both sides of the container, there must be installed gas boxes, each containing 4 SS 8752430 or 4 DIN 13260-2 type oxygen quick-connect sockets and 2 SS 8752430 or 2 DIN 13260-2 type medical compressed air quick-connect sockets. The boxes must be interconnected and the connection must allow for the transfer of gases from one box to another without obstruction. The dimensions of the box must be chosen according to the quick-connect sockets used, but not larger than 800x600x300 mm. | EX
6.8. Equipment
6.8.1. | MUST | All medical devices must comply with Regulation (EU) 2017/745 of European parliament and of the council of 5 April 2017 on medical devices. | EX
6.8.2. | MUST | Monitor-defibrillator – two defibrillators with monitor, recording the physiological parameters of a patient, enabling one to monitor the following parameters of a patient: registration of a 12-lead ECG; registration of a 3- or 5-lead ECG; heart rate measurement with a measuring range of at least from 30 to 300 BPM; pulse oximetry (SpO2) with a measuring range of at least from 1 to 100%, to an accuracy of ± 2%; heart rate with a measuring range of at least from 25 to 240 BPM; monitoring of non-invasive blood pressure (NIBP) with a measuring range of at least from 40 to 270 mmHg; monitoring of invasive blood pressure (IBP) with a measuring range of at least from -30 to 300 mmHg; capnography monitoring using sidestream capnography equipped with Microstream accessories or equivalent technology the Microstream method; the device must provide feedback on the quality of performed chest compressions during resuscitation according to ECR/AHA guidelines; the device must enable round-the-clock registration of administered drugs and trend monitoring of measured parameters on the screen, as well as the printout of trends. The device must be complete with the following tools: a set of cables and electrodes for recording a 12-lead ECG; cable and sensor for monitoring pulse oximetry; different sized cuffs and hoses for non-invasive blood pressure monitoring; invasive blood pressure monitoring kit; capnography sensor; a wall mount; charger; with a protective case and closable accessory pockets (cables, electrodes, sensors, etc.) and a shoulder strap. The device screen must meet the following requirements: a diagonal screen of at least 16 cm, coloured; allow simultaneous monitoring of at least five different measurable signals on the screen; the amplitude of the graph (waveform) must be modifiable. The device must be equipped with an integrated printer, which in turn must meet the following requirements: the width of the printer paper 80 mm; the possibility of ECG lead printing; summary printing of the resuscitation activity; printing alarms and problems; the device must be complete with printer paper (10 pcs). The device must enable a patient’s heart to be affected by an electrical pulse defibrillation or cardioversion if necessary. Defibrillation must comply with the following requirements: application of biphasic electrical pulse; user-friendly, working on principle 1: choose energy – 2: charge – 3: shock; selection of the energy level of the electric pulse at least within the limits of 1 J to 200 J; reaching the energy level of the electric impulse when charging with a full battery 200 J in less than 7 seconds; the function of synchronising the electrical pulse with a patient’s heart rhythm. The device must allow monitoring of the heart rhythm in real-time and, if necessary, to perform extracorporeal electrocardiostimulation (electrocardiostimulator), which shall meet at least the following parameters: the stimulation current shall be changeable from 0 to 140 mA; the stimulation frequency shall be changeable from 30 to 180 beats/min. in increments of 10 beats/min.; enable the use of fixed and, if necessary, demand pacing modes. The device must be equipped with a set of batteries, meeting the following requirements: at least 6 hours of monitoring or transmission of 300 electrical impulses to the human body with an energy of at least 200 J; batteries shall be chargeable from both 230V and 12V power supply systems; automatic charging of batteries in the mains. The device must be equipped with a variable audible and visual alarm system (in compliance with IEC 60601-1-8), signalling at least the following events: battery discharge; high and low heart rate; low SpO2; high and low respiratory rate; high and low non-invasive pressure; high and low invasive pressure; life-threatening cardiac arrhythmia; low or high capnography reading. The device must meet the following general requirements: weight in full assembly up to 5.5 kg; the device must operate in accordance with all the requirements of the Estonian Defence Forces; impact and splash resistant; IEC60529 – IPX5 (water and foreign objects); MIL-STD810G – M514.6; EN ISO 9919 – IEC 60068-2.64; RTCA/DO-160G (helicopters); EN 1789 shock: IEC60068-2-27, 100 g 6 ms half sine; falling EN1789 (0.75 m), IEC 60601 (1 to 2 m). Power supply: 24 V and 230 V. | EX
6.8.3. | MUST | Examination lamp – examination lamp on a movable stand and with an additional wall mount, with a luminous density of 35 000 lx @ 1 m and a CRI 95 LED. | EX
6.8.4. | MUST | Refrigerator – for medicines and solutions, with a capacity of at least 40 L, intended for storing medical preparations and medicines. It must be possible to set the temperature with an accuracy of 0.1 degrees in the range from +2 to +8 °C. The refrigerator must be divided into two sections with one shelf, and the shelf must withstand a load of at least 30 kg. The noise level of the appliance must be less than 50 dB(A) and the weight must be less than 60 kg. | EX
6.8.5. | MUST | Cabinet for medicines and accessories. It must be possible to close it securely for transport. | EX
6.8.6. | MUST | Equipment trolley – two equipment trolleys for mounting monitors referred to in clause 6.8.2..The trolley must also accommodate at least two suspended bags of infusion solution and the attachment of at least two infusion pumps and/or drip counters. The equipment trolley must include a shelf or drawer for accessories. The trolley must be stable and it must be possible to attach the trolley to the container for transportation. | EX
6.8.7. | MUST | Infusion pump – eight infusion pumps with dimensions not exceeding 245 x 90 x 180 mm, easy to use, mobile and easy to clean, weighing not more than 2.0 kg with an integrated fixation clip for the drip stand and a battery. Mains power 230 V, and the device must operate from both the mains and the lithium battery, with a battery operating time of at least 10 h 6 h at an infusion rate of 5 ml/h. The device must allow for the use of different syringes (different types of sizes 5, 10, 20, 30, 50/60 ml, recognised automatically by the infusion pump) and functions: set volume of infusion; specified infusion duration; administration of a bolus; the infusion pump must have an anti-bolus function. The device must have a coloured touch screen, indicating at least the following parameters: infusion rate (ml/h); the amount of infused liquid (at least from 0.1 to 999.9 ml); the time until the infusion is complete/the syringe emptied; occlusion level (settable within the range of approximately 300 to 900 mbar or equivalent from 0 to 1000 mbar and easily modifiable); event trend, at least 1500 per an event. The parameters must be changeable at least in the following range: rate from 0.01 to 1500 ml/h (for 10–50ml syringes); set volume of infusion from 0.1 to 100 ml; infusion duration from 1 min. to 24 hours; a keyboard shortcut must be provided to fill a bolus. The accuracy of the infusion must be at least ± 2% in accordance with the IEC/EN 60601-2-24 standards. The alarms must have both two lights representing two levels and a sound announcing at least the end of the infusion, reaching the set volume of infusion, remaining time until the end of the infusion (a few minutes prior), battery discharge, battery capacity, 230V external mains failure, technical error. The device must be equipped with audible and visual alarms indicating at least infusion completion, reaching the set infusion volume, low battery level, external mains failure, occlusion, and technical errors. Possibility to switch the syringe pumps to the alarm system (nurse call). The instructional text transmitted on the pump screen must be in Estonian or Lithuanian or English. Infrared port, in addition to data transfer, also for inserting new software upgrades. The device must be equipped with an infrared port, USB port, or equivalent interface for data transfer and software upgrades. The syringe plunger must not protrude beyond the syringe pump housing in the case of a filled syringe. Possibility to connect syringe pumps to a special quick-coupled storage compartment with power supply. | EX
6.8.8. | MUST | Device for heating infusion solutions – two devices with a heating temperature of 37 °C and 39 °C 38°C. Start-up time not more than 30 seconds. Capacity of at least 150 ml of liquid per minute (5 °C > 39 °C), at least 1.5 L in succession. Suitable for IV system with diameters of 4 to 5 mm or 6 to 7 mm. Heating profile length 1500 mm, Luer lock connections for attachment of IV- and blood products. Weight 1300 800g; 230V power supply connectors, 3m connection cable, IP X 2 protection rating, rechargeable, Li-ion battery. | EX
6.8.9. | MUST | Field hospital intensive care bed – 20 folding intensive care beds for a field hospital. NSN6530-01-624-7626. | EX
6.8.10. | MUST | Warming cabinet for infusion solutions – with a capacity of at least 20 L. Power consumption of no more than 900 W, internal temperature 39 °C +/- 2 °C. | EX
6.8.11. | MUST | Aspirator – two aspirators with an aspiration capacity of 34 L/min. (± 4 L/min.), with a battery life of at least 60 min. and a charging time not exceeding 4 hours; unladen weight not exceeding 4 kg; weight with a secretion tank not exceeding 5 kg; the noise level less than 60 dB (A); the volume of the secretion tank must be at least 1000 ml. Aspirator power supply of 230V AC and 12V or 24V DC. It must be possible to mount the device on the wall and on a stand. | EX
6.8.12. | MUST | Patient warming device – two electric patient warming blowers with a heating power of at least 460 W and an average power consumption of no more than 800 W. The temperature of the blown air must be adjustable to values of 32 °C, 38 °C, and 43 °C. The device must reach the required temperature within a maximum of 5 minutes. The device must have an overheat alarm. The device must come with consumables for 100 patients. The consumables must include a lower body, upper body, and patient base warming blanket. The device must weigh no more than 8 kg and produce no more than 54 55dB of noise. | EX
6.8.13. | MUST | Oxygen humidifier – two oxygen humidifiers with heating function. The weight should not exceed 2.5 kg. Temperature target values: 31 °C, 34 °C, 37 °C. The heating rate is 10 minutes at 31 °C and 30 minutes at 37 °C. Moisture output is at least 32 mg/l at 37 °C and at least 10 mg/l at 31 °C. The maximum oxygen flow rate is 60 l/min. The device must have an oxygen sensor with an accuracy of ± 4%. The device must come with a disinfection kit. The device must have a device cart or corresponding magnetic wall mounts for the entire device and its accessories. The device must be maintenance-free. | EX
6.8.14. | MUST | Two magnetic wall mounts for infusion pumps/IV drip sets with a capacity of 4 devices. The wall mount must provide power and data connectivity to the devices. | EX
7. STERILIZATION MODULE
7.1. General
7.1.1. | MUST | The container must be capable of being connected to another container of the same type, to allow for the creation of a single system. | EX
7.1.2. | MUST | The walls and floor of the room should be easy to clean and resistant to disinfectant solutions. | EX
7.1.3. | MUST | The container shall have two internal compartments: a room for packaging cleaned equipment and a room for equipment cleaning. | EX
7.1.4. | MUST | The dirty room of the container must contain a stainless steel sink with a worktop. Sink dimensions approx. 400 mm in width, 100 mm in depth and 200 mm in length. The contracting authority may further specify the dimensions and properties of the sink and the worktop during the design phase. The system must include a hot water heater of at least 10 L. | EX
7.1.5. | MUST | There must be a main entrance door at one end of the container and a technical room for equipment at the other end. | EX
7.1.6. | MUST | The container must have two appropriately marked emergency exits, one from each work compartment, for departing the container when the entrance door is blocked. The tenderer shall submit their own solution. | EX | Both container side walls are equipped with emergency escape hatch so that there will be possible to evacuate from both rooms
7.2. Low Voltage Systems
7.2.1. | MUST | The electricity mains of the container is 3-phase 400V 50Hz. The container must be powered from an external supply (grid or generator). In case there is no external supply for the generator the main systems (including lighting) of the container shall be powered by batteries, whereby the batteries shall provide at least an hour at full power. | EX
7.2.2. | MUST | In addition to the 400V and 230V electrical systems, there must be 12V and 24V electrical systems if this is required for the equipment used. The container must be equipped with the required low-voltage/extra low-voltage transformers. | EX
7.2.3. | MUST | The interior power system of the container must be designed to meet the needs of the planned equipment (power outlets types and quantity, IP ratings, lighting, potential equalisation sockets, circuit breakers, protection devices etc) according to all relevant standards and be appropriate for the purpose of the container. | EX
7.3. Extra Low Voltage (ELV) & Data Communication Systems
7.3.1. | MUST | The container must contain RJ45 outlets (eight in total) at workstations and at least two inputs with RJF type sockets in the external feedthrough cabinet, with a possibility to close them with a cap or a stopper. The cables have an endpoint in the ELV & data communication cabinet on the RJ45 panel and they are marked. Cat6 twisted pair cable is used, the colour code of which is T568B. Cables are marked at both ends. | EX
7.3.2. | MUST | An SM fibre optic fibre pair (standard OS-2), with two pairs of fibres in total, shall be run to the external feed panel; the fibres shall be terminated with LC duplex adapters on the communication cabinet panel. On the outer feed panel, a 4ch Stratos HxMA type socket is used; it is closed with a stopper. Plug/socket standard MIL83526/MIL20/MIL21. | MD
7.3.3. | MUST | The ELV & data communication cabinet must be at least 19-inch, which must contain at least a 12-port Cisco (or an analogue programmable on the same principles) switch for setting up tunnels, and a Cisco (or an analogue programmable on the same principles) 4G router with 4G and a GPS antenna on the roof. The cabinet must have a server computer to run the monitoring system, generator and medical equipment software. The cabinet must be provided with proper cooling. | EX
7.4. Illumination
7.4.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | EX
7.4.2. | MUST | Interior lighting must ensure uniform and diffused illumination, meeting the requirements for surgical premises, throughout the room. Illumination must support dimming and DALI control. The illumination of working surfaces at 950 mm from the floor must be at least 1000 lx. | EX
7.4.3. | MUST | The illumination of technical areas and rooms must be at least 200 lx if not stated otherwise. | EX
7.4.4. | MUST | The interior lightning of the container shall be ensured by the generator, external power and battaries, whereby the batteries shall provide at least an hour at full power. | EX
7.5. Ventilation
7.5.1. | MUST | The container must have proper air exchange and it must be equipped with an F9 filter. Preheater and preheated supply air of at least 30 l/s, with preheater adjustable between 5–25 °C. Two-stage filtration, with at least G4 in the first stage and at least F9 in the second stage. Supply air to a clean room, with a pressure difference of at least 5 Pa between the clean and dirty rooms. | MD
7.5.2. | MUST | There must be a pressure differential between the dirty and clean rooms of the container. The pressure in the clean room must be 20 Pa higher compared to the dirty room. Air temperature must be selectable in the range of +18...+25 °C with a tolerance of +/-2 °C at an outside temperature range of -35 to +40 °C, assuming that the room includes up to eight people and the specified equipment. | EX
7.6. Water Supply and Sewerage System
7.6.1. | MUST | The container shall contain a hand-washing station with a stainless steel sink (with rounded corners), measuring at least 675 mm in width, not more then 255 mm in length and at least 200 mm in depth (used orthopedic instruments must fit into the faucet for washing), with a clean water tank for min. 10 L, power of 2kW. | EX
7.6.2. | MUST | Designed and built to be compatible with the interior fittings and purpose of the sterilisation container. The container must support pumping intake water from an external water source. It must include a water purification system for the water supply of steam generators, washers and autoclave cooling. | EX
7.6.3. | MUST | The container shall be equipped with a water purification system that produces water of adequate quality for steam sterilisers. The water purification system shall consist of at least a mechanical filter, a water softener and a reverse osmosis unit. The water purification system must ensure the continuous operation of the washers and autoclaves in the container by providing a 24/7 supply of water. The water purification system must produce water of varying levels of purity (e.g. demineralized and softened water) for the autoclaves and washers. The estimated quality of the input water must be potable water, as determined by the regulation established under Section 85(2) of the Water Act. The water purification system must maintain a water supply reserve for the completion of device cycles in the event of a water connection interruption so that devices can complete their cycles after water is disconnected from the container. The water purification equipment and reservoirs must fit within the dimensions of the container and be usable immediately after unloading. The water purification equipment must be able to be preserved, i.e. all water must be drained from the purification equipment and pipes. Elements that can technically release water must be easily and quickly removable. | EX
7.6.4. | MUST | To ensure a reserve supply of purified water, the container shall have a clean water tank with a volume of at least 50 L. | EX
7.6.5. | MUST | It must be possible to connect the container to an external water mains or to an unpressurised external water tank regardless of the season. | EX
7.6.6. | MUST | It must be possible to collect greywater and contaminated water from the container in an external tank that can be connected to a sewerage network. | EX
7.6.7. | MUST | The container’s water system must not freeze in winter. | EX
7.7. Equipment
7.7.1. | MUST | All medical devices must comply with Regulation (EU) 2017/745 of European parliament and of the council of 5 April 2017 on medical devices. | EX
7.7.2. | MUST | Sterile packaging sealer – Closing speed (output power) – 10m/min; Length of the closing line – limited; Operating pressure at least 10 N; Heat radiation at least 0.1 kJ/s; Closing temperature – 80-220 °C; Closing temperature tolerance ± 2%; Power below 510 Va; Voltage: 230V; Weight less than 25kg; Dimensions 620x260x250 – attachable to the container for transportation. | EX
7.7.3. | MUST | Ultrasonic bath for cleaning instruments – Voltage: 230V; Power: 600W; Current: 3.6 A; Protection class: DIGITEC: IP33; Frequency: 35kHz; Temperature accuracy ± 2.5 °C; Dimensions 325x300x200 – attachable to the container for transportation; Useful volume at least 12l; Drain connection type: G1/2; Weight less than 10kg; Permitted operating temperature 5-40°C. | EX
7.7.4. | MUST | Transport container – two transport trolleys for surgical instruments. | EX
7.7.5. | MUST | Autoclave – the container must include at least two autoclaves with a combined capacity of four STU (Sterilisation Units). Device weight: ≤600kg 520 ; Device height: ≤1900mm 1720mm; Device depth: ≤975 mm 970 mm; Device width: ≤800 mm 690 mm; Chamber volume at least: 160 150 l; Device average power consumption: 7 kW/h; Device maximum instantaneous power consumption: 17 kW; Device maximum cooling water requirement: 100 l/h; Device maximum allowed heat release to the room: 1.9 kW; Surface roughness of the sterilization chamber: Ra 1.25 μm; The device must have an Estonian and English language user interface; The device must have at least an 8 5.7-inch touch screen for operating and selecting sterilization programs. The device screen must display cycle data and graphs. The device must have a printer that prints out the sterilization cycle data. The device must have an RJ45 and USB interface for exporting sterilization cycle data. The system must allow user configuration. The device maintenance menu must provide a clear overview of the device's schematic condition, including temperature, pressure, and valve status. The device must be floor-mounted and also have wheels to ensure necessary workspace around the device during maintenance. The device must be fully automatic, which means that when selecting a program, no other elements need to be configured, including the steam generator and water purification system. The device must be integrated with a steam generator, whose heating elements are made of DIN 1.4571 AISI 316 Ti stainless steel. The chamber of the device must be made of 1.4404 AISI 316 L stainless steel. Both autoclaves must have an emergency stop option in the middle of the sterilization cycle. The system must have a large enough cooling water tank to cool down the sterilizers in case of an emergency stop in any work stage. Autoclave doors must be possible to open with the help of special equipment by a person with appropriate training, even in an emergency stop situation, if the autoclaves cannot be cooled down. The supplier must provide training, instructions, and equipment with the container. | EX
7.7.6. | MUST | Pass-through washers – the container must include at least two pass-through washers with a combined capacity of 400 L. Device noise level ≤ 66 dB; Device weight: <250 kg 180 kg; Device height: <1950 mm 1860 mm; Device depth: <720 mm 630 mm; Device width: <655 mm; Loading height: >800 mm; Water flow rate at least 4–12 L/min 0.5 L/s; Chamber capacity at least: 200 l; The device monitors the amount of detergent in the tank and alerts the user if the amount is insufficient. There is a designated space inside the device for detergent containers, and they can be loaded into the device from the front. The washer has a USB, RJ45, or RS 232 port for exporting wash cycle data. The washer has a control panel from which users can select wash programs, alarms, and view the end time of the wash cycle. Users can create wash programs with a temperature of 90 °C, and existing programs can be modified. The washing rack for loading items into the washer moves on rollers or telescopic rails. The device must have a drainage pump to pump water from the device into the sewage system. The washing chamber can hold at least 8 DIN 1/1 trays at a time, with tray dimensions of 480 x 250 x 50 mm (W x L x H). The offered device is front-loaded and rear-discharged, meaning it has two doors. The lower edge of the washing chamber door is at least 800mm from the floor, ensuring ergonomic loading of laboratory equipment. The washing chamber door opens downward and fixes parallel to the floor to form a loading bridge. The device door has a glass window for viewing the washing chamber. The device door is equipped with automation that prevents the user from opening it during the wash cycle. Each device must come with: 1 instrument basket 1/1 with lid, 2 instrument baskets ½ with lid, 2 instrument baskets ¼, a basket for kidney dishes, a shoe holder, and 4 containers of detergent. The washer must have a channel for steam exhaust that can be connected directly to the room's exhaust system. The device must have an emergency stop function, meaning its operation can be stopped in the middle of a cycle, drained of water, and opened. | EX
7.6.7. | MUST | Folding tables – four folding tables affixed to end walls – two with size 1800x600 mm and two with size 1300x600 mm. | EX
8. MEDICAL STORAGE ROOM (Pharmacy)
8.1. General
8.1.1. | MUST | The doorway of the pharmacy must be designed and constructed in such a way that a connecting corridor with a tent can be attached to the outside of the pharmacy. The container must have a marked emergency exit for departing the container when the entrance door is blocked, the tenderer shall describe their own solution. The container must be capable of being connected to another container of the same type, to allow for the creation of a single system. | EX | The container has designed with a emergency exit hatch/door that can be opened mechanically and from the inside, which can be used in an emergency.
8.1.2. | MUST | The pharmacy storage room must be easy to use, meaning that there may be no partitions or equipment permanently installed in the pathway. | EX
8.1.3. | MUST | The container must have a storage capacity of at least 120 L for medical devices and frozen medicinal products (below -20 °C). | EX
8.1.4. | MUST | The container must have a storage capacity of at least 150 L for medical devices and cool medicinal products (+2 to +8 °C) The storage capability must also ensure that 450ml blood bags can be stored. | EX
8.2. Low Voltage Systems
8.2.1. | MUST | The electricity mains of the container is 3-phase 400V 50Hz. The container must be electrically autonomous, which is ensured, in addition to the possibility of connecting to the external power supply (i.e grid), by a diesel generator with a capacity of at least 8 kW. The fuel supply of the diesel generator must ensure the autonomy of the container for 36 hours. The fuel tank of the container must allow refuelling from outside the container. | EX
8.2.2. | MUST | During the operation of the container's electric generator, it must be possible switch to an external power supply and back without a power interruption. In the case of power interruption, the autonomous electric generator must start and connect to the network within 30 seconds. In case there is no grid connection and the operation of the generator is excluded the main systems (including lighting) of the container shall be powered by batteries, whereby the batteries shall provide at least an hour at full power. Switching with an RLA switch. | EX
8.2.3. | MUST | The standalone generator must start automatically. At startup time, uninterrupted operation of the system must be ensured (e.g. UPS). The container must have a built-in fuel tank for 36 hours of medium load work, and refueling must take place from outside the container. | EX
8.2.4. | MUST | In addition to the 400V and 230V electrical systems, there must be 12V and 24V electrical systems if this is required for the equipment used. The container must be equipped with the required low-voltage/extra low-voltage transformers. | EX
8.2.5. | MUST | The electrical system of the container must be designed and constructed to provide 3-phase 32A 400V and 1-phase 16A 230V outlets to the two high-pressure tents included in the kit. | EX
8.2.6. | MUST | The interior power system of the container must be designed to meet the needs of the planned equipment (power outlets types and quantity, IP ratings, lighting, potential equalisation sockets, circuit breakers, protection devices etc) according to all relevant standards and be appropriate for the purpose of the container. | EX
8.3. Extra Low Voltage (ELV) & Data Communication Systems
8.3.1. | MUST | For workstations (4 pcs), with fibre optic cable (standard OM-3 or OM-4) with two fibres (fibre pair) and terminating in LC duplex adapters, in the ELV & data communication cabinet, fibres are brought to the panel of FO cables, with endpoints in LC adapters and marked at both ends. The container must contain RJ45 outlets (four in total) at workstations and at least two inputs with RJF type sockets in the external feedthrough cabinet, with a possibility to close them with a cap or a stopper. The cables have an endpoint in the ELV & data communication cabinet on the RJ45 panel and they are marked. Cat6 twisted pair cable is used, the colour code of which is T568B. Cables are marked at both ends. | MD
8.3.2. | MUST | An SM fibre optic fibre pair (standard OS-2), with two pairs of fibres in total, shall be run to the external feed panel; the fibres shall be terminated with LC duplex adapters on the ELV & data communication cabinet panel. On the outer feed panel, a Stratos type socket is used; it is closed with a stopper. Plug/socket standard MIL83526/MIL20/MIL21. | EX
8.3.3. | MUST | The ELV & data communication cabinet must be at least 19-inch, which must contain at least 12 ports for setting up a Cisco (or an analogue programmable on the same principles) switch for setting up tunnels, and a Cisco (or an analogue programmable on the same principles) 4G router with 4G and a GPS antenna on the roof. The cabinet must have a server computer for running the monitoring system, generator and medical device software. The cabinet must be provided with proper cooling. | MD
8.4. Illumination
8.4.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | EX
8.4.2. | MUST | Interior lighting shall provide uniform and diffused illumination throughout the room. Illumination must support dimming and DALI control. Illumination on working surfaces at 950 mm from the floor surface must at least be 1000 lx. | EX
8.4.3. | MUST | The illumination of technical areas and rooms must be at least 200 lx if not stated otherwise. | EX
8.4.4. | MUST | The interior lightning of the container shall be ensured by the generator, external power and battaries, whereby the batteries shall provide at least an hour at full power. | EX
8.5. Ventilation
8.5.1. | MUST | The container must exchange air at least 4 times per hour. The air temperature must be selectable in the range of +18...+25 °C with fluctuations of +/-2 °C at an outside temperature of -35 to +40 °C in a room for up to eight people with the described equipment. Air intake with a heat exchanger of at least 30 l/s. | EX
8.6. Water Supply and Sanitation – None
8.7. Medical Gases and Oxygen – None
8.8. Equipment – None
9. LOGISTICS CONTAINER
9.1. General
9.1.1. | MUST | The container, designed and built on a steel frame, must be integrated with the hook-lift system (hook-loop in front and rails under the bottom). | EX
9.1.2. | MUST | In the front of the container there must be a technical room for the refrigeration machine, fuel tank and other necessities. | EX
9.1.3. | MUST | The flooring must be grooved aluminium, slip-resistant. | EX
9.1.4. | MUST | The walls must be fitted with a combi-track to which both the load belt and the transport posts can be attached. | EX
9.1.5. | MUST | The floor of the container must have fastening loops (10–12 pcs), equally divided, shall be specified during the design stage. | EX
9.1.6. | MUST | The container must maintain an internal temperature of +2 to +18 oC. | EX
9.1.7. | MUST | It must be possible to maintain the temperature in the Medical Logistics Container in accordance with the Medicinal Products Act in force in Estonia and/or Lithuania. | EX
9.1.8. | MUST | The Medical Logistics Container has a door at one end with a minimum size of 1400 x 2100 mm. | EX
9.2. Low Voltage Systems
9.2.1. | MUST | The electricity mains of the container is 3-phase 400V 50Hz. The container must be powered from an external supply (grid or generator). In case there is no external supply for the generator the main systems (including lighting) of the container shall be powered by batteries, whereby the batteries shall provide at least an hour at full power. The power supply must be ensured for all medical functions and container technical systems that are necessary for opening and closing the container without an external power source or starting the generator. The battery system must withstand at least four such opening and closing cycles, as well as charging and discharging cycles. In addition, the power supply must ensure the operation of the container's automatic and/or alarm system for 48 hours. The battery system must ensure uninterrupted operation of the container's diesel heating equipment (Webasto) for at least 48 hours. | EX
9.2.2. | MUST | The low voltage cabinet of the container must have a 3-phase 400V 32A input and 2x 1-phase 16A outputs, sockets are in the cabinet. | EX
9.2.3. | MUST | Container must be equipped with a refrigeration machine Carrier Supra 850 (diesel + electric power supply) or equivalent. Fuelling must be brought to the container's exterior surface. | EX
9.2.4. | MUST | The interior power system of the container must be designed to meet the needs of the planned equipment (power outlets types and quantity, IP ratings, lighting, potential equalisation sockets, circuit breakers, protection devices etc) according to all relevant standards and be appropriate for the purpose of the container. | EX
9.3. Extra Low Voltage (ELV) & Data Communication Systems – None
9.4. Illumination
9.4.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | EX
9.4.2. | MUST | Interior lighting shall provide uniform and diffused illumination throughout the room. Illumination must support dimming. The illumination of working surfaces at 950 mm from the floor must be at least 200 lx. | EX
9.4.3. | MUST | The interior lightning of the container shall be ensured by the generator, external power and battaries, whereby the batteries shall provide at least an hour at full power. | EX
9.5. Ventilation – None
9.6. Water Supply and Sanitation – None
9.7. Medical Gases and Oxygen – None
9.8. Equipment – None
10. RESOURCE MODULE 1
10.1. General
10.1.1. | MUST | The container must provide the necessary resources for the use of the Mobile Modular Hospital solution. | EX
10.1.2. | MUST | Electricity generation of at least 175 kW (exact capacity to be determined during the design phase); The container must have a built-in fuel tank for 36 hours of medium load work, and refueling must take place from outside the container. | MD
10.1.3. | MUST | The electric generator must be the same as the paraghaph 11 in the resource container. | EX
10.1.4. | MUST | Production of oxygen with a purity of more than 90% for a capacity of at least 80 L per minute, between 4.5–5.5 bar, the elements of oxygen production must be removable from the container. | EX
10.1.5. | MUST | Bottling of oxygen: fill two 10L balloons at once in 70 minutes to a pressure of 200 bar. Oxygen distribution sockets must be provided in accordance with the number of oxygen production elements and the DIN 13260-2 or SS 8752430 standard. It must also be possible to supply oxygen to other containers by supplying them with the oxygen produced. Container has to be supplyed with an oxygen hose of at least 120 m in length and an inside diameter of 8 mm. In addition, hose ends (60 + 60 mounting clips) must be provided as spare parts. A gas panel with 6 SS 8752430 type oxygen quick-connect sockets must be installed on the side of the container. The dimensions of the panel must be chosen according to the quick-connectors used, but must not be larger than 800x600x300 mm. | EX
10.1.6. | MUST | The need for a fire barrier between the oxygen production system and the electricity generator will be determined during the design phase. | EX
10.1.7. | MUST | The noise of the container must not exceed 70 db, measured at a distance of 7 m, at any direction. | EX
10.1.8. | MUST | The container shall be equipped with a sufficient number of extension cables to power the two systems, the quantities and types of outlets for the electrical distribution shield of the container will be determined during the design phase. | EX
10.1.9. | MUST | The container must have an in/out penetration cabinet for the electrical system and the network connection located on the outside of the container. | EX
10.1.10. | MUST | It must be possible to use the container on a trailer and mounted on ground level. | EX
10.1.11. | MUST | External walls – 2mm steel sheet | EX
10.1.12. | MUST | Insulation – Stone wool-filled three-layer sandwich panels. The outer and inner surfaces of the sandwich panels shall be a 0.5 mm thick steel sheet with a polyester surface coating (S280+Z275 PE) glued to the core. The materials used must be safe for humans and the environment. | EX
10.1.13. | MUST | Floors – expected service loads for floors are at minimum 3.0 kN/m². The equipment room of the container must be equipped with a collecting tray to prevent oils and fuels from leaking to the outside environment. | EX
10.1.14. | MUST | Ceiling roof – construction analogous to walls, design snow load on roofs shall be at least 2.0 kN/m². The roof must bear movement on the roof and the clearing of snow. | EX
10.1.15. | MUST | Insulation – The generator room must be separated from other rooms by a noise-cancelling fire wall with a fire resistance of min. EI30 and surrounded by noise and fire barriers. | EX
10.2. Low Voltage Systems
10.2.1. | MUST | The electricity mains of the container is 3-phase 400V 50Hz with TN-S wiring. | EX
10.2.2. | MUST | The container must be equipped with an alternative power supply input, the purpose of which is to connect the container to an external power source in order to ensure the entire R2 system with electricity from the existing power supply. The inlet for the alternate power supply must be with a 20m cable and the solution for the inlet shall be provided by the tenderer. The kit must include the appropriate cable. | EX
10.2.3. | MUST | The container and all its outlets must be equipped with a separate circuit breaker and, if necessary, a fault current circuit limiter, in accordance with the legislation in force in Estonia to ensure maximum safety of personnel and equipment. | EX
10.2.4. | MUST | The generator must be capable of producing electricity for 36 hours at 75% load without the additional need for refuelling. Fuelling the generator must take place from the outside. | EX
10.3. Extra Low Voltage (ELV) & Data Communication Systems
10.3.1. | MUST | Four RJ-45 wall sockets (places will be specified during the design phase) must be installed in the intranet. In the external penetration cabinet, two ELV inputs are brought into the cable cabinet. Outdoor penetrations are RJF type slots that can be closed with a cap or a stopper. The cables have an endpoint in a ELV & data communication cabinet on the RJ45 panel and they are marked. Cat6 twisted pair cable is used, the colour code of which is T568B. Cables are marked at both ends. | EX
10.3.2. | MUST | The ELV & data communication cabinet must be at least 19-inch, which must contain at least a 12-port Cisco (or an analogue programmable on the same principles) switch for setting up tunnels, and a Cisco (or an analogue programmable on the same principles) 4G router with 4G and a GPS antenna on the roof. The cabinet must have a server computer to run the monitoring system, generator and medical equipment software. The cabinet must be provided with proper cooling. | EX
10.4. Illumination
10.4.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | EX
10.4.2. | MUST | Interior lighting shall provide uniform and diffused illumination throughout the room. Illumination must support dimming. The illumination of working surfaces at 950 mm from the floor must be at least 200 lx. | EX
10.4.3. | MUST | The interior lightning of the container shall be ensured by the generator, external power and battaries, whereby the batteries shall provide at least an hour at full power. In addition, illumination must be available in the generator room. Electrical supply for the generator room illumination must be ELV to ensure safety for working with the generator. | EX
10.5. Heating, Cooling and Ventilation
10.5.1. | MUST | The ventilation in the generator room must be designed and constructed in accordance with the requirements and must ensure the smooth operation of the generator throughout the temperature range. The whole system must be able to operate with the doors closed. | EX
10.5.2. | MUST | The air intake, both in storage and in operation, must be filtered in such a way as to avoid indoor pollution, in such a way that it is safe for people to be in there. | EX
10.5.3. | MUST | When storing, ensure an internal temperature of at least +10 °C and air humidity of less than 50% throughout the outdoor temperature range. | EX
10.6. Water Supply and Sanitation – None
10.7. Medical Gases and Oxygen – None
10.8. Equipment – None
11. RESOURCE MODULE 2
11.1 General
11.1.1. | MUST | The need for a fire barrier between the water production system and the electricity generator will be determined during the design phase. | EX
11.1.2. | MUST | Electricity generation of at least 175kW (exact capacity to be determined during the design phase). The container must have a built-in fuel tank for 36 hours of medium-load work, refueling must be done from outside the container. The electrical generator must be the same as the one in the resource container 1. | MD
11.1.3. | MUST | It must be possible to use the container on a trailer and mounted on ground level. | EX
11.1.4. | MUST | External walls – 2mm steel sheet. | EX
11.1.5. | MUST | Insulation – Stone wool-filled three-layer sandwich panels. The outer and inner surfaces of the sandwich panels shall be a 0.5 mm thick steel sheet with a polyester surface coating (S280+Z275 PE) glued to the core. The materials used must be safe for humans and the environment. | EX
11.1.6. | MUST | Floors – calculated service loads for floors are at minimum 3.0 kN/m². The equipment room of the container must be equipped with a collecting tray to prevent oils and fuels from leaking to the outside environment. | EX
11.1.7. | MUST | The noise of the container must not exceed 70 db, measured at a distance of 7 m, the generator room must be surrounded with noise and fire barriers. | EX
11.1.8. | MUST | The container shall be equipped with a sufficient number of extension cables to power the two systems, the quantities and types of outlets for the electrical distribution shield of the container will be determined during the design phase. | EX
11.2. Low Voltage Systems
11.2.1 | MUST | The electricity mains of the container is 3-phase 400V 50Hz with TN-S wiring. | EX
11.2.2. | MUST | The container must be equipped with an alternative power supply input, the purpose of which is to connect the container to an external power source in order to ensure the entire R2 system with electricity from the existing power supply. The inlet for the alternate power supply must be with a 20m cable and the solution for the inlet shall be provided by the tenderer. The kit must include the appropriate cable. | EX
11.2.3. | MUST | The container and all its outlets must be equipped with a separate circuit breaker and, if necessary, a fault current circuit limiter, in accordance with the legislation in force in Estonia to ensure maximum safety of personnel and equipment. | EX
11.2.4. | MUST | The generator must be capable of producing electricity for 36 hours at 75% load without the additional need for refuelling. Fuelling the generator must take place from the outside. | MD
11.3. Extra Low Voltage (ELV) & Data Communication Systems
11.3.1. | MUST | Four RJ45 sockets will be mounted on the walls (location to be specified during the design phase), and two low-current inputs with RJF type sockets in the external penetration cabinet, which can be sealed with a cap or a stopper.The cables have an endpoint in a ELV & data communication cabinet on the RJ45 panel and they are marked. Cat6 twisted pair cable is used, the colour code of which is T568B. Cables are marked at both ends. | EX
11.3.2. | MUST | The container must contain at least two ELV inputs. The ELV & data communication cabinet must be at least 19-inch, which must contain at least a 12-port Cisco (or an analogue programmable on the same principles) switch for setting up tunnels, and a Cisco (or an analogue programmable on the same principles) 4G router with 4G and a GPS antenna on the roof. The cabinet must have a server computer to run the monitoring system, generator and medical equipment software. The cabinet must be provided with proper cooling. | EX
11.4. Illumination
11.4.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | EX
11.4.2. | MUST | Interior lighting shall provide uniform and diffused illumination throughout the room. Illumination must support dimming. The illumination of working surfaces at 950 mm from the floor must be at least 200 lx. | EX
11.4.3. | MUST | The interior lightning of the container shall be ensured by the generator, external power and battaries, whereby the batteries shall provide at least an hour at full power. In addition, illumination must be available in the generator room. Electrical supply for the generator room illumination must be ELV to ensure safety for working with the generator. | EX
11.5. Heating, Cooling and Ventilation
11.5.1. | MUST | The ventilation in the generator room must be designed and constructed in accordance with the requirements and must ensure the smooth operation of the generator throughout the temperature range. The whole system must be able to operate with the doors closed. | EX
11.5.2. | MUST | The air intake, both in storage and in operation, must be filtered in such a way as to avoid indoor pollution. | EX
11.5.3. | MUST | When storing the container, ensure an internal temperature of at least +10 °C and air humidity of less than 50% throughout the outdoor temperature range. | EX
11.6. Water Supply and Sanitation – None
11.7. Medical Gases and Oxygen – None
11.8. Equipment – None
11.9. Production of Domestic (Potable) Water
11.9.1. | MUST | The container must be able to produce all the water needed for the field hospital from uncontaminated surface water (river, lake). The container must be able to purify at least 500 L of purified water per hour. The purified water quality must comply with the regulation established on the basis of § 85 (2) of the Water Act. | EX
11.9.2. | MUST | The container shall be capable of drawing ground water from the water body with a suction pump, and shall be equipped with an electric submersible pump in case the water source does not allow suction pumping (distance or difference in height). The kit must include the necessary hoses. | EX
11.9.3. | MUST | The container shall be able to draw inlet water from a water tower or a reservoir, etc. | EX
11.9.4. | MUST | The water must be treated mechanically, chemically and with ultraviolet radiation in order to obtain drinking water that meets the requirements. The resource of the cleaning equipment must be such that it can be automatically regenerated. It must be possible to distribute the water to other containers/containers with water connections. The types of water connections will be specified during the design phase. The water treatment equipment must be possible to preserve, i.e., all the water must be drained from the treatment equipment and pipelines. Elements from which it is not technically possible to drain water must be easily and quickly removable. Double spare parts kits must be provided for all cost elements requiring maintenance (e.g., filters, membranes, filter media). | EX
11.9.5. | MUST | The container shall allow for the storage of at least 500 L of water for domestic use within containers and a total of 5500 L of water. Outside the container, used water may be stored in a container that does not retain its shape. | EX
11.9.6. | MUST | The outlet for clean water shall be pressurised with pressure provided by an automatic water dispenser or a pressure booster pump of at least 1.5 kW capacity. It must be possible to pump the output water to the tanks, to the consumers of water, via the water network of the field hospital. It shall also be possible to pump water to the water collectors or tanks from the tanks included in the kit. | EX
12. SANITARY MODULE
12.1. General
12.1.1. | MUST | The container must be possible to attach to a container truck using twistlock fastenings. | EX
12.1.2. | MUST | The container must accommodate at least 5 toilets and 3 shower areas. The toilets must have sinks for handwashing as well as soap and paper towel dispensers. | EX
12.1.3. | MUST | The container may be box-based and access to them may be provided from the exterior sides of the container. The technical room may be provided as a separate room. The solution must also consider providing private changing facilities. The exterior doors must be made of metal. | EX
12.1.4. | MUST | The container must be equipped with a ladder or stairs that allow users to access the container when it is located on a container truck or Mack trailer. | EX
12.1.5. | MUST | All sanitary equipment used in the container, including toilets and sinks, must be made of stainless steel. | EX
12.1.6. | MUST | All equipment, small inventory, and fixtures must be attached to the container in such a way that they do not damage or compromise the contents of the container during transportation. | EX
12.2. Low Voltage Systems
12.2.1. | MUST | The electricity mains of the container is 3-phase 400V 50Hz. The container must be powered from an external supply (grid or generator). In case there is no external supply for the generator the main systems (including lighting) of the container shall be powered by batteries, whereby the batteries shall provide at least an hour at full power. | EX
12.2.2 | MUST | The interior power system of the container must be designed to meet the needs of the planned equipment (power outlets types and quantity, IP ratings, lighting, potential equalisation sockets, circuit breakers, protection devices etc) according to all relevant standards and be appropriate for the purpose of the container. The container must be designed and construct according to Estonian Standard EVS-HD 60364-7-701 or Lithuanian Standard LST HD 60364-7-701: Requirements for special installations or locations – Locations containing a bath or shower | EX
12.3. lllumination
12.3.1. | MUST | The illumination will be based on LED lighting and the solution must also include exterior lighting for illuminating all four sides of the container. To ensure overall reliability of the illumination, different types of light sources must be kept to a minimum. | EX
12.3.2. | MUST | Interior lighting shall provide uniform and diffused illumination throughout the room. Illumination must support dimming. The illumination of working surfaces at 950 mm from the floor must be at least 200 lx. | EX
12.4. Heating, Cooling and Ventilation
12.4.1. | MUST | The container must be equipped with heating and ventilation equipment in the general area to ensure the required indoor climate. | EX
12.5. Water and Sewerage System
12.5.1. | MUST | To provide hot water, the container must have a boiler. The size of the boiler must ensure the supply of hot water for 3 showers and washbasins. The recommended minimum boiler capacity is 400 liters and the heating element should be at least 12 kW. Ideally, the boiler's power should be 2-stage (step switching). | EX
12.5.2. | MUST | The container must be possible to connect to an external water and sewage system or temporary tanks. The water supply pipes must be dimensioned in such a way as to ensure a flow of water for 5 toilets, 3 showers, washbasins, and washing machines, with a minimum diameter of 40 mm. | EX
12.5.3. | MUST | If access to the showers is through the container, a drain trap must be installed in front of the shower cubicles. In addition, curtains or doors must be installed for the showers in such a way that water does not spill onto the floor and the curtains or doors are not damaged during transportation. | EX
12.6. Equipment
12.6.1. | MUST | The container must accommodate and contain 2 washing machines and one dryer (9-13 kg). When selecting the equipment, higher usage intensity than that of a typical household user must be taken into account. The washing machines must be conservable, meaning that if necessary, water must be able to be drained from them. | EX
12.6.2. | MUST | Hooks must be installed next to each shower and toilet cubicle for storing clothes in the container. For external doors, the hooks must be on the inside. | EX
13.POWER GENERATORS
13.1. General
13.1.1. | MUST | A three-phase diesel generator (hereafter ‘generator’), a fuel tank, a switchboard, extension cables and other necessary additional equipment shall be permanently installed inside the container (hereafter ‘container’). | EX
13.1.2. | MUST | The container shall have at least one emergency stop button (protected against mechanical damage and with a protection class IP67) on the outer side of the container. The exact location of the emergency stop button shall be coordinated during construction. | EX
13.1.3. | MUST | It must be possible to operate a running generator with the container doors closed, even if the consumers are connected to the switchboard via extension cables. | EX
13.1.4. | MUST | Additional equipment installed inside the container must not interfere with the operation of the generator and must be easily accessible for maintenance or repair of the equipment. | EX
13.1.5. | MUST | All parts of the metal construction of the container must be bonded to the structure thereof (electrical bonding in electrical terms). | EX
13.1.6. | MUST | All installation cables must consist of fiber copper cable. | EX
13.1.7. | MUST | There must be a powder fire extinguisher and CO2 extinguisher fastened inside the container (in total, 2 pcs per container). The fire extinguisher must be at least 6kg 5kg ABC CO2 and secured so as to be easily accessible and withstand vibration when transporting the container. | EX
13.1.8. | MUST | The completed system must be designed to operate under the following conditions: | EX
13.1.8.1. | MUST | weather conditions according to STANAG 4370 AECP 200. A3 Intermediate Dry or equivalent and STANAG 4370 AECP 200. C1 Intermediate Cold or equivalent; | EX
13.1.8.2. | MUST | under inclination (the maximum inclination shall be determined during the project); | EX
13.1.8.3. | MUST | where appropriate, the use of a diesel burner as a preheater shall be permitted, provided that the starter is battery-operated and can be pre-programmed to warm up the diesel engine for at least 30 to 120 minutes. | EX
13.2. Requirements for the Generator I
13.2.1. | MUST | The generator must be certified according to standard EVS-EN ISO8528 (in its entirety) – to ensure prime power without a time limit. | EX
13.2.2. | MUST | Active power shall be determined during the design phase with the aim of enabling all consumers that are part of the system to be fed. | EX
13.2.3. | MUST | The quality class of the electricity network must be certified for at least G3, according to standard EVS-EN ISO 8528 (in its entirety). | EX
13.2.4. | MUST | Network frequency 50 Hz. | EX
13.2.5. | MUST | Output phase voltage of 230 V and line voltage of 400 V. | EX
13.2.6. | MUST | The generator must have a brushless type alternator. | EX
13.2.7. | MUST | The output wiring of the electrical installation of the generator must be in the TN-S system (separated by the protective conductor), the neutral (N) and protective earth (PE) conductors permanently connected (PEN) at a single point on the supply side of the electrical installation (switchboard or generator). | EX
13.2.8. | MUST | The part of the space allocated to the generator of the container must be equipped with a pre-programmable preheater that runs on diesel fuel. | EX
13.2.9. | MUST | Diesel engine and preheater (without additional modification) must be designed to run on the following fuels: EVS-EN 590:2022+NA:2022/ LST EN 590:2022/NA:2022; NATO STANAG 7090 type F63 and F54. | EX
13.2.10. | MUST | The capacity of the generator’s fuel tank must be such as to allow uninterrupted operation at full power for at least 36 hours without intermediate refuelling. The tank must have an electronic and mechanical fuel level indicator, which must be independent, a sediment separator and a coarse filter made of a glass body. | EX
13.2.11. | MUST | The location of the generator’s fuel tank must allow simple refuelling and visual inspection of the fuel level indicator. The tenderer shall describe their solution in detail. | EX | Genset fuel tank level indicator is located so it can be visually checked from outside the container and fuel level is indicated on genset control panel
13.2.12. | MUST | The diesel engine shall be liquid-cooled with the exhaust system located on the end of the container. | EX
13.2.13. | MUST | The diesel engine must be fitted with a tap or hose system on the bottom of the crankcase, allowing for the easy collection of oil during an oil change. | EX
13.2.14. | MUST | The interval between regular service intervals (or replacement of liquids and filters) must be at least 250 engine hours (except in the run-in period or other specific requirement by the manufacturer). | EX
13.2.15. | MUST | The rotational speed of an internal combustion engine must be 1000 or 1500 rpm. | EX
13.2.16. | MUST | The connector of the ground cable must be bolted and fitted with a wing nut located at two points on the outer side of the container. The grounding cable must have a butterfly clip fastener. Removal of the butterfly nut/bolt must be prohibited. | EX
13.2.17. | MUST | The grounding pole (must be positioned as attached to the container and as close as possible to the ground connection) must be previously calculated and tested by the manufacturer (confirmed by the corresponding certificate). The grounding pole must be made of stainless steel or equivalent. | EX
13.2.18. | MUST | It must also be possible to refuel the generator safely with a running engine from both canisters and refuelling pistols at service stations in EU countries. | EX
13.2.19. | MUST | The fuel tank must be fitted with an air vent constructed in such a way as to prevent any leakage of fuel when the container is being lifted or transported (electronic valve or similar ignition system). | EX
13.2.20. | MUST | The acoustic sound pressure level of the generator with doors closed may not exceed 70 dB(A), measured at a distance of 7m from the container, running continuously at 75% of the prime power rating. | EX
13.2.21. | MUST | The generator must include automatic charging equipment for the starter battery, used in case of storage or emergency preparedness (supplied from 230 V 50 Hz external network and available 24/7/365). When stored without charging, the starting battery must be guaranteed to function for a minimum of 6 months without any additional switches. | EX
13.2.22. | MUST | The generator must be equipped with an automatic coolant heater, used in case of storage or emergency preparedness (supplied from 230 V 50 Hz external network and available 24/7/365). Heating must be ensured through the same connection as charging. The heating must be separately controllable (the cord is plugged in, the battery is charging, but the heating is not working/working). If the hospital is fully staffed, then the "standby" generator should be in quick readiness without an additional switch, meaning the block is warm and the battery is charging. | EX
13.2.23. | MUST | The generator unit must not be installed in a double housing (i.e. double housing/container around alternator and engine). | EX
13.2.24. | MUST | The generator control panel must display at least: | EX
13.2.24.1. | MUST | effective phase-neutral voltage and voltage between each phase; | EX
13.2.24.2. | MUST | network frequency (Hz); | EX
13.2.24.3. | MUST | power consumption (active, reactive and full power); | EX
13.2.24.4. | MUST | counter of operating hours; | EX
13.2.24.5. | MUST | oil temperature and oil pressure; | MD
13.2.24.6. | MUST | temperature of coolant; | EX
13.2.24.7. | MUST | an indicator for specifying and configuring maintenance times (it must be possible to set the maintenance time when a certain number of operating hours have been reached, as indicated with the sounding of a corresponding audio alarm and a text alert displayed on the control panel, indicating the need to arrange the necessary maintenance); | EX
13.2.24.8. | MUST | brief descriptions of all possible error codes (longer code descriptions must be included in the user manual of the generator); | EX
13.2.25. | MUST | The generator control panel must have at least an English menu. | EX
13.2.26. | MUST | The generator control panel must have a Start/Stop button (preheating must be automatic). | EX
13.2.27. | MUST | The generator control panel must have an adjustable a backlight enabling the readability of the data both during the use of ceiling lighting and in daylight. | EX
13.2.28. | MUST | The generator must be equipped with at least the following protective devices: | EX
13.2.28.1. | MUST | disconnect switch for starter battery (disconnects both the power and control circuits of the entire generator), when disconnecting the battery, the maintenance charger function must be preserved. | EX
13.2.28.2. | MUST | automatic voltage regulator (AVR); | EX
13.2.28.3. | MUST | automatic stop in case of low oil pressure level; | MD
13.2.28.4. | MUST | prewarning (for on-time fuelling) in case of low fuel level; | EX
13.2.28.5. | MUST | automatic stop, a warning (with sufficient reference to reducing the load) must precede it in case of electrical overload; | EX
13.2.28.6. | MUST | automatic stop at least in case of too high oil and coolant temperature; | EX
13.2.28.7. | MUST | automated main circuit breaker for automatic disconnection of all loads. | EX
13.2.29. | MUST | The generator control panel must sound an alarm and display an alert message if any operating parameter (including the need to arrange maintenance) deviates from the stated range values required by the manufacturer. | EX
13.3. Requirements for the Generator II
13.3.1. | MUST | The metal switchboard must be mounted inside the container in a vibration-proof manner (the exact locations and quantities shall be determined during the design phase). | EX
13.3.2. | MUST | The switchboard must be made of non-corrosive metal for the protection of surfaces or equivalent. | EX
13.3.3. | MUST | The switchboard must be dimensioned according to the maximum permitted power of output sockets. | EX
13.3.4. | MUST | The switchboard must have a protection class of at least IP65. | EX
13.3.5. | MUST | All outputs and the input must have a protection class of IP67 and, on the outer side of the container, be protected by a weatherproof lockable keyless door. | EX
13.3.6. | MUST | The residual-current device must be of class “A”. | EX
13.3.7. | MUST | The switchboard must be protected from overvoltage with a surge protector of class “C” (DIN). | EX
13.3.8. | MUST | An earthing system of the switchboard must be TN-S. | EX
13.3.9. | MUST | The switchboard must come with an integrated multimeter for load measurement (at least per each phase: amperes (A), volts (V), kilowatts (kW), kilovolt-amperes (kVa), frequency (Hz) and network Cos (φ)). The multimeter must have a backlight, the intensity of which shall be adjustable in a range enabling the readability of the data both during the use of ceiling lighting and in daylight. | EX
13.3.10. | MUST | Power cables must be installed in a non-corrosive metal cable duct to protect cables from mechanical damage. | EX
13.3.11. | MUST | The switchboard must be equipped with outputs and an input with residual-current circuit breakers. The quantity and types of electrical outputs shall be determined during the design phase based on the number and nature of consumers). | EX
13.3.12. | MUST | All markings on the inputs and outputs of the electrical distribution board shall be clearly and durably marked in accordance with the as-built drawings. | EX
13.3.13. | MUST | The generator must be accompanied by guides both in electronic (.pdf or MS Office compatible file format, allowing search in text and copying of text in text format from the document) and in hard-copy format, with the following levels of detail: | EX
13.3.14. | MUST | a user-level operation and maintenance manual (made of water-proof material) in Estonian/Lithuanian and English that must include a checklist for performing functional checks to identify the causes of potential faults; | EX
13.3.15. | MUST | a maintenance engineer’s level maintenance and repair manual at least in English. | EX
13.3.16. | MUST | The generator must be supplied with a list of delivered parts, as well as respective drawings, in Estonian and in English in both electronic and hard-copy format. The list of delivered parts must cover all elements included in the generator delivery. | EX
13.3.17. | MUST | The generator must be supplied with a list of delivered consumables and spare parts, in Estonian/ Lithuanian and in English in both electronic and hard-copy format. The list of delivered consumables and spare parts must include descriptions and products codes (as well as NSN codes, if available) of the consumables and spare parts used in maintenance and repair. | EX
13.3.18. | MUST | The generator must be accompanied by at least an English wiring diagram in electronic (.pdf or .dwg or MS Office file format compatible) and in hard-copy format with each cable and connection marked (for the whole container and its associated equipment). The wiring diagram must contain all electrical elements of the container assembly (including the wiring diagram of the generator, the distribution network, the lighting equipment, the generator control systems, etc.). A qualified electrician must be able to use the wiring diagram for troubleshooting and assembly work, if necessary. | EX
13.3.19. | MUST | The generator set must be accompanied by a brief safety manual in Estonian/Lithuanian, in A5 or A4 waterproof format, for the general user. Texts and, where appropriate, images or drawings of the safety instructions shall be agreed upon with the end-user at the time of performance of the contract. | EX
13.3.20. | MUST | The file format of electronic manuals must be .pdf or a format compatible with MS Office, supporting a text search function and the possibility to copy text from the document. The files must be stored on CD/DVD discs; the Contracting Authority will not accept documents delivered on a memory stick, hard drive or another data medium. | EX
13.3.21. | MUST | All user manuals in the hard-copy format must be packaged in a moisture- and impact-resistant plastic or aluminium case that meets the following requirements: moisture and dust resistant, UV-resistant; hinges of stainless steel; can withstand temperatures from -35 °C to +50 °C; lid with lockable keyless clips; carrying handle, external colour black (RAL 9005), military green (RAL 6031) or green (RAL 6025). | EX
13.3.22. | MUST | The Tenderer must provide, at least in digital format, a drawing, in at least English, of the entire container and the equipment inside thereof. The drawing must contain at least the detailed dimensions of the whole structure, the details of reconstruction and the exact location of the equipment. The drawing must indicate the designations of the devices, together with the product codes. | MD
13.3.23. | MUST | Any manuals and catalogues in Estonian and English must be prepared in accordance with EVS-EN IEC/IEEE 82079-1:2020 or any manuals and catalogues in Lithuanian and English must be prepared in accordance with LST EN IEC/IEEE 82079-1:2020. | EX
13.3.24. | MUST | After the delivery of the goods, the Tenderer must carry out a type training of the user and maintenance technician level of the generator in Tallinn for at least 8 academic hours for persons designated by the Contracting Authority (max 10 persons) at the place of business thereof. | EX
13.3.25. |  | The type training must include at least the following: | EX
13.3.26. | MUST | operation of the generator (general presentation, control panel readouts and meaning of alarm signals/warnings, identification and use of protective devices, locations and use of sockets, identification and use of control equipment, adding fuel); | EX
13.3.27. | MUST | The type training must include at least the following maintenance of the generator (change of fluids, filters, flushing of the fuel tank, internal and external cleaning of the housing, checking of the operating mode, checking of the technical condition, main adjustments, checking of the fastenings and shock absorbers); | EX
13.3.28. | MUST | typical mistakes in generator use, common error messages, troubleshooting, elimination of faults; | EX
13.3.29. | MUST | overview of main consumables and spare parts (descriptions and order codes). | EX
13.3.30. | MUST | The exact time and place of the training shall be agreed upon reception of the goods. | EX
13.3.31. | MUST | The training shall be provided in the Estonian and English language. | EX
13.3.32. | MUST | The Tenderer must issue to all participants in the training a corresponding training certificate, certifying the right of further training of the persons who have completed the training. | EX
13.3.33. | MUST | The Tenderer must provide each trainee with training and distribution materials in electronic or in hard-copy format. | EX
13.3.34. | MUST | Prior to the training, the Tenderer must communicate to the Contracting Authority the agenda of the training and the name(s) of the person(s) conducting the training. | EX
13.3.35. | MUST | The generator must be supplied with all necessary accessories, consumables, spare parts and tools required for performing the first three regular maintenances. Non-moisture-resistant equipment (filters, seals, tools, etc., where appropriate) must be packaged in a moisture- and impact-resistant plastic or aluminium case that meets the following requirements: moisture and dust resistant, UV-resistant; hinges of stainless steel; can withstand temperatures from -35 °C to +50 °C; lid with lockable keyless clips; carrying handle, external colour black (RAL 9005), military green (RAL 6031) or forest green (RAL 6025). | EX
13.3.36. | MUST | Warranty repair works on all the equipment offered must be carried out by the Tenderer or their authorized representative. | EX
13.3.37. | MUST | The Tenderer must indicate the contact details of the agencies in Estonia responsible for the sale of consumables and spare parts for the equipment, and maintenance and repair work thereof. | EX
13.3.38. | MUST | The Tenderer must ensure the availability of all consumables and spare parts of the equipment supplied in Estonia or Lithuania (as per the buyer) for at least 10 years after the expiry of the warranty period. | EX
13.3.39. | MUST | EVS-HD 60364-5-551:2010/AC:2011/ LST HD 60364-5-551:2010/AC:2010 Electrical installations of buildings. Part 5-55: Selection and erection of electrical equipment. Other equipment. Clause 551: Low-voltage generating sets; | EX
13.3.40. | MUST | EVS-HD 60364-1:2008/ LST HD 60364-1:2008 Low-voltage electrical installations. Part 1: Fundamental principles, assessment of general characteristics, definitions; | EX
13.3.41. | MUST | EVS-HD 60364-4-41:2007/ LST HD 60364-4-41:2017Low-voltage electrical installations. Part 4-41: Protection for safety. Protection against electric shock; | EX
13.3.42. | MUST | EVS-HD 60364-4-42:2011+A1:2015/LST HD 60364-4-42:2011+A1:2015 Low voltage electrical installations. Part 4-42: Protection for safety. Protection against thermal effects; | EX
13.3.43. | MUST | EVS-HD 60364-4-43:2010/LST HD 60364-4-43:2010 Low-voltage electrical installations. Part 4-43: Protection for safety. Protection against overcurrent; | EX
13.3.44. | MUST | EVS-HD 60364-5-534:2016/LST HD 60364-5-534:2016 Low-voltage electrical installations. Part 5-53: Selection and erection of electrical equipment. Isolation, switching and control. Clause 534: Devices for protection against overvoltage; | EX
13.3.45. | MUST | EVS-EN 12601:2010 Reciprocating internal combustion engine driven generating sets. Safety; | EX
13.3.46. | MUST | the generator must comply with standard IEC 60034; | EX
13.3.47. | MUST | the distribution network cable must comply with standard IEC-60228 Class 5 (H07RN-F); | EX
13.3.48. | MUST | all installation must be carried out in accordance with current EU standards and norms; | EX
13.3.49. | MUST | all products offered must bear the CE marking. | EX
13.3.50. | MUST | the container must comply with the requirements of the International Convention for Safe Containers and bear the CSC marking. | EX
13.3.52. | MUST | The generator must meet the requirements of ISO 8528. | EX
13.3.53. | MUST | Marking plates must me made of durable material (aluminium or equivalent), which is riveted to the housing of the product. Gluing is not allowed. | EX
13.3.54. | MUST | The marking must be applied to the marking label by laser engraving or other equivalent technology. | EX
13.3.55. | MUST | The housing of electrical equipment must include an electrical warning sign. | EX
13.3.55. | MUST | For moving the container, there must be a safety instruction on the outer side thereof, indicating the maximum permitted angle of inclination of the container when unloading, loading and transporting the container. | EX
13.3.56. | MUST | The container marking must also contain the CSC safety approval data. | DV
13.3.57. | MUST | In addition to the previous the marking label shall also indicate: | EX
13.3.58. | MUST | Designation: container; Commodity code: 8-digit commodity code (Estonian Defence Forces storage code, the data of which shall be communicated to the Tenderer during the design phase) | EX
13.3.59. | MUST | In addition to the manufacturer’s marking, the generator must be marked with the following information: | EX
13.3.62. | MUST | Product designation; Data, characterising the switchboard according to the EU requirements. | EX
13.3.63. | MUST | The generators to be installed must be the same throughout the system. | EX
13.3.64. | MUST | A diesel generator (hereafter ‘generator’), a fuel tank, a switchboard, extension cables and other necessary additional equipment shall be permanently installed inside the container (hereafter ‘container’). | EX
13.3.65. | MUST | The container shall have an emergency stop button (protected against mechanical damage and with a protection class IP67) on the outer side of the container. The exact location of the emergency stop button shall be coordinated during construction. | EX
13.3.66. | MUST | It must be possible to operate a running generator with the container doors closed, even if the consumers are connected to the switchboard via extension cables. | EX
13.3.67. | MUST | Additional equipment installed inside the container must not interfere with the operation of the generator and must be easily accessible for maintenance or repair of the equipment. | EX
13.3.68. | MUST | All parts of the metal construction of the container must be bonded to the structure thereof (electrical bonding in electrical terms). | EX
13.3.69. | MUST | All installation cables must consist of fiber copper cable. | EX
13.3.70. | MUST | There must be a powder fire extinguisher and CO2 extinguisher fastened inside the container (in total, 2 pcs per container). The fire extinguisher must be secured so as to be easily accessible and withstand vibration when transporting the container. | EX
13.4. The Completed System Must Be Designed to Operate Under the Following Conditions:
13.4.1. | MUST | weather conditions according to STANAG 4370 AECP 200. A3 Intermediate Dry or equivalent and STANAG 4370 AECP 200. C1 Intermediate Cold or equivalent; | EX
13.4.2. | MUST | under the inclination (the maximum inclination shall be determined during the project); | EX
13.4.3. | MUST | where appropriate, the use of a diesel burner as a preheater shall be permitted, provided that the starter is battery-operated and can be pre-programmed to warm up the diesel engine for at least 30 to 120 minutes; | EX
13.4.4. | MUST | The generator must be certified according to standard EVS-EN ISO8528 (in its entirety) – to ensure prime power without a time limit. | EX
13.4.5. | MUST | Active power shall be determined during the design phase with the aim of enabling all consumers that are part of the system to be fed. | EX
13.4.6. | MUST | The quality class of the electricity network must be certified for at least G2, according to standard EVS-EN ISO 8528 (in its entirety). | EX
13.4.7. | MUST | Network frequency 50 Hz. | EX
13.4.8. | MUST | The output phase voltage of 230 V and line voltage of 400 V (whether three-phase or single-phase, shall be determined during the design phase). | EX
13.4.9. | MUST | The generator must have a brushless type alternator (shall be determined during the design phase). | EX
13.4.10. | MUST | The output wiring of the electrical installation of the generator must be in the TN-S system (separated by the protective conductor), the neutral (N) and protective earth (PE) conductors permanently connected (PEN) at a single point on the supply side of the electrical installation (switchboard or generator). | EX
13.4.11. | MUST | The part of the space allocated to the generator of the container must be equipped with a pre-programmable preheater that runs on diesel fuel (shall be determined during the design phase). | EX
13.4.12. | MUST | Diesel engine and preheater (without additional modification) must be designed to run on the following fuels: EVS-EN 590:2013+NA:2014; NATO STANAG 7090 type F63 and F54. | EX
13.4.12. | MUST | The capacity of the generator’s fuel tank must be such as to allow uninterrupted operation at full power for at least 36 hours without intermediate refuelling. The tank must have an electronic and mechanical fuel level indicator, a sediment separator and a coarse filter made of a glass body. | EX
13.4.13. | MUST | The location of the generator’s fuel tank must allow simple refuelling and visual inspection of the fuel level indicator. | EX
13.4.14. | MUST | The diesel engine shall be liquid-cooled with the exhaust system located on the end of the container. | EX
13.4.15. | MUST | The diesel engine must be fitted with a tap or hose system on the bottom of the crankcase, allowing for the easy collection of oil during an oil change. | EX
13.4.16. | MUST | The interval between regular service intervals (or replacement of liquids and filters) must be at least 250 engine hours (except in the run-in period or other specific requirement by the manufacturer). | EX
13.4.17. | MUST | The connector of the ground cable must be bolted and fitted with a wing nut located at two points on the outer side of the container. | EX
13.4.18. | MUST | The grounding pole (must be positioned as attached to the container and as close as possible to the ground connection) must be previously calculated and tested by the manufacturer (confirmed by the corresponding certificate). The grounding pole must be of stainless steel. | EX
13.4.19. | MUST | It must also be possible to refuel the generator safely with a running engine from both canisters and refuelling pistols at service stations in EU countries. | EX
13.4.20. | MUST | The fuel tank must be fitted with an air vent constructed in such a way as to prevent any leakage of fuel when the container is being lifted or transported (electronic valve or similar ignition system) (shall be determined during the design phase). | EX
13.4.21. | MUST | The acoustic sound pressure level of the generator with doors closed may not exceed 70 dB(A), measured at a distance of 7 m from the container, running continuously at 75% of the prime power rating. | EX
13.4.22. | MUST | The generator must include automatic charging equipment for the starter battery, used in case of storage or emergency preparedness (supplied from 230 V 50 Hz external network and available 24/7/365). | EX
13.4.23. | MUST | The generator must be equipped with an automatic coolant heater, used in case of storage or emergency preparedness (supplied from 230 V 50 Hz external network and available 24/7/365). | EX
13.4.24. | MUST | The generator unit must not be installed in a double housing (i.e. double housing/ container around alternator and engine). | EX
13.4.25. | MUST | The generator control panel must display at least: | EX
13.4.26. | MUST | effective phase-neutral voltage and voltage between phases (unit V); | EX
13.4.27. | MUST | network frequency (Hz); | EX
13.4.28. | MUST | power consumption (active, reactive and full power); | EX
13.4.29. | MUST | counter of operating hours; | EX
13.4.30. | MUST | oil temperature and oil pressure; | MD
13.4.31. | MUST | temperature of coolant; | EX
13.4.32. | MUST | an indicator for specifying and configuring maintenance times (it must be possible to set the maintenance time when a certain number of operating hours have been reached, as indicated with the sounding of a corresponding audio alarm and a text alert displayed on the control panel, indicating the need to arrange the necessary maintenance); | EX
13.4.33. | MUST | brief descriptions of all possible error codes (longer code descriptions must be included in the user manual of the generator); | EX
13.4.34. | MUST | The generator control panel must be duplicated in the room part of the medical container itself and allow to control the generator. | EX
13.4.35. | MUST | The generator control panel must have at least an English menu. | EX
13.4.36. | MUST | The generator control panel must have a Start/Stop button. | EX
13.4.37. | MUST | The generator control panel must have an adjustable backlight enabling the readability of the data both during the use of ceiling lighting and in daylight. | EX
13.4.38. | MUST | The generator must be equipped with at least the following protective devices: | EX
13.4.39. | MUST | disconnect switch for starter battery (disconnects both the power and control circuits of the entire generator); | EX
13.4.40. | MUST | automatic voltage regulator (AVR); | EX
13.4.41. | MUST | automatic stop in case of low oil level; | MD
13.4.42. | MUST | automatic stop in case of low fuel level; | EX
13.4.43. | MUST | automatic stop in case of electrical overload; | EX
13.4.44. | MUST | automatic stop at least in case of too high oil and coolant temperature; | MD
13.4.45. | MUST | automated main circuit breaker for automatic disconnection of all loads. | EX
13.4.46. | MUST | The generator control panel must sound an alarm and display an alert message if any operating parameter (including the need to arrange maintenance) deviates from the stated range values required by the manufacturer. | EX
13.4.47. | MUST | The metal switchboard must be mounted inside the container in a vibration-proof manner (the exact locations and quantities shall be determined during the design phase); | EX
13.4.48. | MUST | The switchboard must be made of metal and treated with an anti-corrosion agent for the protection of surfaces. | EX
13.4.49. | MUST | The switchboard must be dimensioned according to the maximum permitted power of output sockets. | EX
13.4.50. | MUST | The switchboard must have a protection class of at least IP65. | EX
13.4.51. | MUST | All outputs and the input must have a protection class of IP67 and, on the outer side of the container, be protected by a weatherproof lockable keyless door. | EX
13.4.52. | MUST | The residual-current device must be of class “A”. | EX
13.4.53. | MUST | The switchboard must be protected from overvoltage with a surge protector of class “C” (DIN). | EX
13.4.54. | MUST | An earthing system of the switchboard must be TN-S. | EX
13.4.55. | MUST | The switchboard must come with an integrated multimeter for load measurement (at least per each phase: amperes (A), volts (V), kilowatts (kW), kilovolt-amperes (kVa), frequency (Hz) and network Cos (φ)). The multimeter must have a backlight, the intensity of which shall be adjustable in a range enabling the readability of the data both during the use of ceiling lighting and in daylight. | EX
13.4.56. | MUST | Power cables must be installed in a non-corrosive metal cable duct to protect cables from mechanical damage. | EX
13.4.57. | MUST | The switchboard must be equipped with outputs and an input with residual-current circuit breakers. The quantity and types of electrical outputs shall be determined during the design phase based on the number and nature of consumers). | EX
13.4.58. | MUST | All markings on the inputs and outputs of the electrical distribution board shall be clearly and durably marked in accordance with the as-built drawings. | EX
13.4.59. | MUST | The generator must be accompanied by guides both in electronic (.pdf or MS Office compatible file format, allowing search in text and copying of text in text format from the document) and in hard-copy format, with the following levels of detail: | EX
13.4.60. | MUST | a user-level operation and maintenance manual in Estonian/Lithuanian and English that must include a checklist for performing functional checks to identify the causes of potential faults; | EX
13.4.61. | MUST | a maintenance engineer’s level maintenance and repair manual at least in English. | EX
13.4.62. | MUST | The generator must be supplied with a list of delivered parts, as well as respective drawings, in Estonian and in English in both electronic and hard-copy format. The list of delivered parts must cover all elements included in the generator delivery. | EX
13.4.63. | MUST | The generator must be supplied with a list of delivered consumables and spare parts, in Estonian and in English in both electronic and hard-copy format. The list of delivered consumables and spare parts must include descriptions and products codes (as well as NSN codes, if available) of the consumables and spare parts used in maintenance and repair. | EX
13.4.64. | MUST | The generator must be accompanied by at least an English wiring diagram in electronic (.pdf or .dwg or MS Office file format compatible) and in hard-copy format with each cable and connection marked (for the whole container and its associated equipment). The wiring diagram must contain all electrical elements of the container assembly (including the wiring diagram of the generator, the distribution network, the lighting equipment, the generator control systems, etc.). A qualified electrician must be able to use the wiring diagram for troubleshooting and assembly work, if necessary. | EX
13.4.65. | MUST | The generator set must be accompanied by a brief safety manual in Estonian/ Lithuanian, in A5 or A4 waterproof format, for the general user. Texts and, where appropriate, images or drawings of the safety instructions shall be agreed upon with the end-user at the time of performance of the contract. | EX
13.4.66. | MUST | The file format of electronic manuals must be .pdf or a format compatible with MS Office, supporting a text search function and the possibility to copy text from the document. The files must be stored on CD/DVD discs; the Contracting Authority will not accept documents delivered on a memory stick, hard drive or another data medium. | EX
13.4.67. | MUST | All user manuals in the hard-copy format must be packaged in a moisture- and impact-resistant plastic or aluminium case that meets the following requirements: moisture and dust resistant, UV-resistant; hinges of stainless steel; can withstand temperatures from -35 °C to +50 °C; lid with lockable keyless clips; carrying handle, external colour black (RAL 9005), military green (RAL 6031) or forest green (RAL 6025). | EX
13.4.68. | MUST | The Tenderer must provide, at least in digital format, a drawing, in at least English, of the entire container and the equipment inside thereof. The drawing must contain at least the detailed dimensions of the whole structure, the details of reconstruction and the exact location of the equipment. The drawing must indicate the designations of the devices, together with the product codes. | EX
13.4.69. | MUST | Any manuals and catalogues in Estonian and English must be prepared in accordance with EVS-EN IEC/IEEE 82079-1:2020 or any manuals and catalogues in Lithuanian and English must be prepared in accordance with LST EN IEC/IEEE 82079-1:2020. | EX
13.4.70. | MUST | After the delivery of the goods, the Tenderer must carry out a type training of the user and maintenance technician level of the generator in Tallinn for at least 8 academic hours for persons designated by the Contracting Authority (max 10 persons) at the place of business thereof. | EX
13.4.71. | MUST | The type training must include at least the following: | EX
13.4.71.1. | MUST | operation of the generator (general presentation, control panel readouts and meaning of alarm signals/warnings, identification and use of protective devices, locations and use of sockets, identification and use of control equipment, adding fuel); | EX
13.4.71.2. | MUST | The type training must include at least the following maintenance of the generator (change of fluids, filters, flushing of the fuel tank, internal and external cleaning of the housing, checking of the operating mode, checking of the technical condition, main adjustments, checking of the fastenings and shock absorbers); | EX
13.4.71.3. | MUST | typical mistakes in generator use, common error messages, troubleshooting, elimination of faults; | EX
13.4.71.4. | MUST | overview of main consumables and spare parts (descriptions and order codes). | EX
13.4.72. | MUST | The exact time and place of the training shall be agreed upon receipt of the goods. | EX
13.4.73. | MUST | The training should be provided in the Estonian and English language. | EX
13.4.74. | MUST | The Tenderer must issue to all participants in the training a corresponding training certificate, certifying the right of further training of the persons who have completed the training. | EX
13.4.75. | MUST | The Tenderer must provide each trainee with training and distribution materials in electronic or in hard-copy format. | EX
13.4.76. | MUST | Prior to the training, the Tenderer must communicate to the Contracting Authority the agenda of the training and the name(s) of the person(s) conducting the training. | EX
13.4.77. | MUST | The generator must be supplied with all necessary accessories, consumables, spare parts and tools required for performing the first three regular maintenances. Non-moisture-resistant equipment (filters, seals, tools, etc., where appropriate) must be packaged in a moisture- and impact-resistant plastic or aluminium case that meets the following requirements: moisture and dust resistant, UV-resistant; hinges of stainless steel; can withstand temperatures from -35 °C to +50 °C; lid with lockable keyless clips; carrying handle, external colour black (RAL 9005), military green (RAL 6031) or forest green (RAL 6025). | EX
13.4.78. | MUST | Warranty repair works on all the equipment offered must be carried out by the Tenderer or its authorized representative. | EX
13.4.79. | MUST | The Tenderer must indicate the contact details of the agencies in Estonia/Lithuania responsible for the sale of consumables and spare parts for the equipment, and maintenance and repair work thereof. | EX
13.4.80. | MUST | The Tenderer must guarantee the availability of spare parts for the working life of the generators, i.e. 10 years. | EX
13.4.81. | MUST | EVS-HD 60364-5-551:2010/AC:2011/ LST HD 60364-5-551:2010/AC:2010 Electrical installations of buildings. Part 5-55: Selection and erection of electrical equipment. Other equipment. Clause 551: Low-voltage generating sets; | EX
13.4.82. | MUST | EVS-HD 60364-1:2008/LST HD 60364-1:2008 Low-voltage electrical installations. Part 1: Fundamental principles, assessment of general characteristics, definitions; | EX
13.4.83. | MUST | EVS-HD 60364-4-41:2007/LST HD 60364-4-41:2017 Low-voltage electrical installations. Part 4-41: Protection for safety. Protection against electric shock; | EX
13.4.84. | MUST | EVS-HD 60364-4-42:2011+A1:2015/LST HD 60364-4-42:2011+A1:2015 Low voltage electrical installations. Part 4-42: Protection for safety. Protection against thermal effects; | EX
13.4.85. | MUST | EVS-HD 60364-4-43:2010/LST HD 60364-4-43:2010 Low-voltage electrical installations. Part 4-43: Protection for safety. Protection against overcurrent; | EX
13.4.86. | MUST | EVS-HD 60364-5-534:2016/LST HD 60364-5-534:2016 Low-voltage electrical installations. Part 5-53: Selection and erection of electrical equipment. Isolation, switching and control. Clause 534: Devices for protection against overvoltage; | EX
13.4.87. | MUST | EVS-EN 12601:2010 Reciprocating internal combustion engine driven generating sets. Safety; | EX
13.4.88. | MUST | the generator must comply with standard IEC 60034; | EX
13.4.89. | MUST | the distribution network cable must comply with standard IEC-60228 Class 5 (H07RN-F); | EX
13.4.90. | MUST | all installation must be carried out in accordance with current EU standards and norms; | EX
13.4.91. | MUST | all products offered must bear the CE marking. | EX
13.4.92. | MUST | the container must comply with the requirements of the International Convention for Safe Containers and bear the CSC marking. | EX
13.4.93. | MUST | The generator must meet the requirements of ISO 8528. | EX
13.4.94. | MUST | Marking plates must be made of durable material (aluminium or equivalent), which is riveted to the housing of the product. Gluing is not allowed. | EX
13.4.95. | MUST | The marking must be applied to the marking label by laser engraving or other equivalent technology. | EX
13.4.96. | MUST | The housing of electrical equipment must include an electrical warning sign. | EX
13.4.97. | MUST | For moving the container, there must be a safety instruction on the outer side thereof, indicating the maximum permitted angle of inclination of the container when unloading, loading and transporting the container. | EX
13.4.98. | MUST | The container marking must contain also the CSC safety approval data. | EX
13.4.99. | MUST | In addition to the previous the marking label shall also indicate: | EX
13.4.100. | MUST | Designation: container; Commodity code: 8-digit commodity code (Estonian Defence Forces storage code, the data of which shall be communicated to the Tenderer during the construction) | EX
13.4.101. | MUST | In addition to the manufacturer’s marking, the generator must be marked with the following information: | EX
13.4.102. | MUST | Product designation; Data, characterising the switchboard according to the EU requirements. | EX
ASUTUSESISESEKS KASUTAMISEKS
Riigi Kaitseinvesteeringute Keskus
Märge tehtud: 08.06.2026
Kehtib kuni: 08.06.2036
Alus: AvTS § 35 lg 1 p-d 4, 17
Note on C113: Silvia Põder: Pole vajalik
Note on C115: Silvia Põder: Pole vaja Lõppkasutaja poolt
Note on C116: Silvia Põder: Ei kohaldu kõigile. Lisa juurde kommentaar et rakendub moodulitel, millel on seda vaja.
Note on C169: Silvia Põder: Ei kohaldu
Note on C488: Silvia Põder: ??
Note on C626: Silvia Põder: Pole vaja, ei kohaldu
Note on C653: Silvia Põder: Lisa juurde pulber kustuti ja kaal 5 kg peale.